## Impedance (Mag&Phase)
Probe is fabricated with a 12 cm length of micro-coax attached to the probe head.
Impedance measurements have been taken with and without fixture compensation for this probe lead.
When measured with a 50resistor termination, the 12 cm micro-coax appears as a 15 nH series inductor.
Probe design
Lithium Nibate Composite
350 um pitch with 30 um kerfs
Piezo thickness 68 um
12 um paralene matching layer
E-Solder conductive epoxy backing
Un-compensated |  |  | Compensated
Frequency | Magnitude | Phase | Frequency | Magnitude | Phase
10000000 | 436.8099 | -88.4626 | 10000000 | 595.9691 | -88.3951
10140000 | 430.3123 | -88.4325 | 10140000 | 586.8555 | -88.3774
10280000 | 423.8976 | -88.4 | 10280000 | 578.1362 | -88.3114
10420000 | 417.7551 | -88.3526 | 10420000 | 569.4023 | -88.2624
10560000 | 411.7175 | -88.3079 | 10560000 | 561.0793 | -88.2084
10700000 | 405.7401 | -88.2465 | 10700000 | 552.8577 | -88.142
10840000 | 400.0009 | -88.1856 | 10840000 | 544.7706 | -88.05
10980000 | 394.528 | -88.0685 | 10980000 | 537.1713 | -87.943
11120000 | 389.4296 | -87.9673 | 11120000 | 530.1995 | -87.7833
11260000 | 384.5468 | -87.9196 | 11260000 | 523.9082 | -87.7196
11400000 | 379.8187 | -87.8765 | 11400000 | 517.444 | -87.7073
11540000 | 375.1052 | -87.8833 | 11540000 | 511.0542 | -87.6784
11680000 | 370.3259 | -87.8624 | 11680000 | 504.468 | -87.6987
11820000 | 365.5201 | -87.8631 | 11820000 | 497.8142 | -87.6784
11960000 | 360.6943 | -87.8342 | 11960000 | 491.1941 | -87.6717
12100000 | 355.9302 | -87.8117 | 12100000 | 484.6005 | -87.624
12240000 | 351.2633 | -87.7605 | 12240000 | 478.1196 | -87.5646
12380000 | 346.7483 | -87.7058 | 12380000 | 471.8274 | -87.5048
12520000 | 342.3117 | -87.6366 | 12520000 | 465.5908 | -87.443
12660000 | 337.9914 | -87.5652 | 12660000 | 459.5885 | -87.3603
12800000 | 333.7474 | -87.4849 | 12800000 | 453.6731 | -87.2741
12940000 | 329.5973 | -87.4168 | 12940000 | 447.7858 | -87.1978
13080000 | 325.4382 | -87.3497 | 13080000 | 441.9489 | -87.0944
13220000 | 321.4063 | -87.2658 | 13220000 | 436.1755 | -86.9559
13360000 | 317.4309 | -87.1517 | 13360000 | 430.5668 | -86.8696
13500000 | 313.4962 | -87.0536 | 13500000 | 425.0443 | -86.7128
13640000 | 309.7008 | -86.9388 | 13640000 | 419.6322 | -86.5884
13780000 | 305.9015 | -86.8029 | 13780000 | 414.2328 | -86.4302
13920000 | 302.1801 | -86.6877 | 13920000 | 409.0171 | -86.2713
14060000 | 298.4651 | -86.5385 | 14060000 | 403.7622 | -86.0927
14200000 | 294.8672 | -86.3938 | 14200000 | 398.5842 | -85.9073
14340000 | 291.1462 | -86.2126 | 14340000 | 393.4056 | -85.683
14480000 | 287.555 | -86.0173 | 14480000 | 388.2343 | -85.4333
14620000 | 284.0202 | -85.7913 | 14620000 | 383.1748 | -85.1573
14760000 | 280.5955 | -85.5105 | 14760000 | 378.1954 | -84.7933
14900000 | 277.246 | -85.1835 | 14900000 | 373.5891 | -84.3928
15040000 | 274.0436 | -84.9102 | 15040000 | 368.8777 | -84.0089
15180000 | 270.965 | -84.5311 | 15180000 | 364.5219 | -83.5168
15320000 | 268.175 | -84.1156 | 15320000 | 360.8337 | -82.9834
15460000 | 265.6105 | -83.6621 | 15460000 | 357.1337 | -82.4051
15600000 | 263.235 | -83.1658 | 15600000 | 354.0341 | -81.7791
15740000 | 261.6196 | -82.6094 | 15740000 | 351.8262 | -81.0636
15880000 | 260.4265 | -82.0938 | 15880000 | 350.668 | -80.3959
16020000 | 259.4828 | -81.683 | 16020000 | 349.6354 | -79.8741
16160000 | 258.5061 | -81.258 | 16160000 | 349.0861 | -79.2799
16300000 | 257.8253 | -80.8316 | 16300000 | 348.7126 | -78.6981
16440000 | 257.4955 | -80.4113 | 16440000 | 349.1814 | -78.1575
16580000 | 257.7501 | -80.0298 | 16580000 | 350.2382 | -77.6196
16720000 | 258.3886 | -79.7297 | 16720000 | 352.1997 | -77.1629
16860000 | 259.3582 | -79.5694 | 16860000 | 354.6452 | -76.8593
17000000 | 260.6596 | -79.4822 | 17000000 | 358.1541 | -76.7424
17140000 | 261.8457 | -79.6279 | 17140000 | 361.3696 | -76.8911
17280000 | 262.7058 | -79.9168 | 17280000 | 364.4773 | -77.2157
17420000 | 263.1392 | -80.3159 | 17420000 | 366.4484 | -77.7135
17560000 | 262.9522 | -80.7511 | 17560000 | 367.4181 | -78.2736
17700000 | 262.3171 | -81.2347 | 17700000 | 367.2305 | -78.8689
17840000 | 261.151 | -81.6898 | 17840000 | 366.6068 | -79.4924
17980000 | 259.5158 | -82.1433 | 17980000 | 364.8081 | -80.0708
18120000 | 257.6496 | -82.5227 | 18120000 | 362.6955 | -80.67
18260000 | 255.7676 | -82.8725 | 18260000 | 360.2993 | -81.1047
18400000 | 253.6685 | -83.2365 | 18400000 | 357.6289 | -81.5476
18540000 | 251.3509 | -83.5574 | 18540000 | 354.6544 | -81.9963
18680000 | 249.1429 | -83.8141 | 18680000 | 351.3749 | -82.3813
18820000 | 246.8108 | -84.0978 | 18820000 | 347.8396 | -82.7841
18960000 | 244.2173 | -84.3183 | 18960000 | 344.1529 | -83.0952
19100000 | 241.7831 | -84.5055 | 19100000 | 340.4517 | -83.3921
19240000 | 239.1708 | -84.6476 | 19240000 | 336.5876 | -83.6345
19380000 | 236.5174 | -84.7634 | 19380000 | 332.669 | -83.7776
19520000 | 234.0473 | -84.8286 | 19520000 | 328.9682 | -83.8479
19660000 | 231.4911 | -84.8565 | 19660000 | 325.1718 | -83.9088
19800000 | 229.2027 | -84.849 | 19800000 | 321.3408 | -83.9111
19940000 | 226.934 | -84.8363 | 19940000 | 318.0512 | -83.9075
20080000 | 224.8106 | -84.7813 | 20080000 | 314.8807 | -83.8651
20220000 | 222.6646 | -84.7515 | 20220000 | 311.8273 | -83.8233
20360000 | 220.5402 | -84.6816 | 20360000 | 308.7126 | -83.7654
20500000 | 218.4357 | -84.5498 | 20500000 | 305.7684 | -83.6315
20640000 | 216.6781 | -84.4161 | 20640000 | 303.1025 | -83.4595
20780000 | 214.9574 | -84.2667 | 20780000 | 300.4825 | -83.2687
20920000 | 213.3735 | -84.1179 | 20920000 | 298.4272 | -83.0442
21060000 | 211.8672 | -83.9456 | 21060000 | 296.6151 | -82.8698
21200000 | 210.7794 | -83.7796 | 21200000 | 295.1149 | -82.6735
21340000 | 209.6678 | -83.6375 | 21340000 | 293.5992 | -82.5036
21480000 | 208.7621 | -83.517 | 21480000 | 292.843 | -82.3254
21620000 | 207.9199 | -83.468 | 21620000 | 291.9257 | -82.26
21760000 | 207.19 | -83.4465 | 21760000 | 291.2297 | -82.2549
21900000 | 206.4119 | -83.4671 | 21900000 | 290.517 | -82.3009
22040000 | 205.6183 | -83.5725 | 22040000 | 289.6365 | -82.4143
22180000 | 204.6011 | -83.6621 | 22180000 | 288.554 | -82.5558
22320000 | 203.6005 | -83.8193 | 22320000 | 287.2836 | -82.7578
22460000 | 202.2629 | -83.9774 | 22460000 | 285.538 | -82.9452
22600000 | 200.8874 | -84.0835 | 22600000 | 283.7351 | -83.1677
22740000 | 199.4581 | -84.2072 | 22740000 | 281.6056 | -83.267
22880000 | 197.8916 | -84.2551 | 22880000 | 279.2213 | -83.3772
23020000 | 196.3493 | -84.3305 | 23020000 | 277.195 | -83.4512
23160000 | 194.9641 | -84.3408 | 23160000 | 275.1033 | -83.5108
23300000 | 193.5005 | -84.3947 | 23300000 | 273.1961 | -83.5641
23440000 | 192.0824 | -84.4176 | 23440000 | 271.0021 | -83.6028
23580000 | 190.5783 | -84.4411 | 23580000 | 268.9173 | -83.6959
23720000 | 189.1914 | -84.4366 | 23720000 | 267.0315 | -83.7147
23860000 | 187.8255 | -84.4537 | 23860000 | 264.8098 | -83.6982
24000000 | 186.2653 | -84.4213 | 24000000 | 262.8834 | -83.6536
24140000 | 184.8826 | -84.407 | 24140000 | 260.6025 | -83.6799
24280000 | 183.5033 | -84.3669 | 24280000 | 258.6687 | -83.6245
24420000 | 182.2474 | -84.2945 | 24420000 | 256.6877 | -83.5704
24560000 | 180.8125 | -84.2638 | 24560000 | 254.8978 | -83.5142
24700000 | 179.5967 | -84.1962 | 24700000 | 253.1188 | -83.4172
24840000 | 178.4324 | -84.1068 | 24840000 | 251.3808 | -83.3627
24980000 | 177.2223 | -84.072 | 24980000 | 249.7097 | -83.3088
25120000 | 176.0164 | -83.9959 | 25120000 | 247.9781 | -83.2337
25260000 | 174.8095 | -83.9301 | 25260000 | 246.1942 | -83.2028
25400000 | 173.5713 | -83.8471 | 25400000 | 244.5665 | -83.0533
25540000 | 172.5059 | -83.7452 | 25540000 | 242.9787 | -82.9552
25680000 | 171.4521 | -83.6249 | 25680000 | 241.5199 | -82.8285
25820000 | 170.4044 | -83.5114 | 25820000 | 240.1208 | -82.6498
25960000 | 169.4827 | -83.4085 | 25960000 | 238.8224 | -82.5609
26100000 | 168.6132 | -83.3159 | 26100000 | 237.7084 | -82.4386
26240000 | 167.8073 | -83.2417 | 26240000 | 236.5552 | -82.3409
26380000 | 166.9834 | -83.1944 | 26380000 | 235.5044 | -82.3274
26520000 | 166.0826 | -83.1678 | 26520000 | 234.2923 | -82.3037
26660000 | 165.0096 | -83.1415 | 26660000 | 232.8243 | -82.2467
26800000 | 164.0471 | -83.0709 | 26800000 | 231.5432 | -82.1435
26940000 | 163.11 | -82.9756 | 26940000 | 230.2303 | -82.0405
27080000 | 162.1551 | -82.8671 | 27080000 | 228.951 | -81.9234
27220000 | 161.3061 | -82.7936 | 27220000 | 227.7259 | -81.7922
27360000 | 160.5143 | -82.6686 | 27360000 | 226.6692 | -81.6441
27500000 | 159.7499 | -82.5621 | 27500000 | 225.7926 | -81.5671
27640000 | 159.0441 | -82.4822 | 27640000 | 224.8223 | -81.4407
27780000 | 158.2751 | -82.3949 | 27780000 | 224.0677 | -81.353
27920000 | 157.5898 | -82.3231 | 27920000 | 223.0493 | -81.2304
28060000 | 156.9305 | -82.2544 | 28060000 | 222.2682 | -81.1636
28200000 | 156.3169 | -82.1858 | 28200000 | 221.5913 | -81.0825
28340000 | 155.7157 | -82.1454 | 28340000 | 220.8849 | -81.0412
28480000 | 155.0713 | -82.1255 | 28480000 | 220.085 | -81.0172
28620000 | 154.4156 | -82.098 | 28620000 | 219.3988 | -81.013
28760000 | 153.7625 | -82.0993 | 28760000 | 218.4344 | -80.9849
28900000 | 153.0882 | -82.067 | 28900000 | 217.6647 | -80.9915
29040000 | 152.3832 | -82.0586 | 29040000 | 216.7935 | -80.9782
29180000 | 151.6846 | -82.0749 | 29180000 | 215.8662 | -80.9645
29320000 | 150.9286 | -82.0356 | 29320000 | 214.9154 | -80.956
29460000 | 150.2295 | -82.0485 | 29460000 | 214.01 | -80.9525
29600000 | 149.4733 | -82.0282 | 29600000 | 213.0072 | -80.9548
29740000 | 148.7496 | -82.0094 | 29740000 | 211.9595 | -80.9487
29880000 | 147.9381 | -81.9709 | 29880000 | 210.8775 | -80.9302
30020000 | 147.1851 | -81.9406 | 30020000 | 209.953 | -80.8882
30160000 | 146.4141 | -81.9028 | 30160000 | 208.9027 | -80.8296
30300000 | 145.7168 | -81.8469 | 30300000 | 207.8959 | -80.7932
30440000 | 145.0215 | -81.7696 | 30440000 | 207.0356 | -80.7074
30580000 | 144.3482 | -81.7041 | 30580000 | 206.1139 | -80.6283
30720000 | 143.7671 | -81.6491 | 30720000 | 205.399 | -80.5887
30860000 | 143.1509 | -81.5916 | 30860000 | 204.6875 | -80.5188
31000000 | 142.64 | -81.5536 | 31000000 | 203.8939 | -80.4881
31140000 | 142.0596 | -81.5441 | 31140000 | 203.2852 | -80.4814
31280000 | 141.4123 | -81.5332 | 31280000 | 202.5312 | -80.4709
31420000 | 140.8744 | -81.535 | 31420000 | 201.7584 | -80.4691
31560000 | 140.1745 | -81.5555 | 31560000 | 200.9433 | -80.5011
31700000 | 139.5293 | -81.5513 | 31700000 | 200.0806 | -80.5005
31840000 | 138.8105 | -81.5345 | 31840000 | 199.1804 | -80.5248
31980000 | 138.135 | -81.528 | 31980000 | 198.1311 | -80.5166
32120000 | 137.4432 | -81.5144 | 32120000 | 197.2574 | -80.5185
32260000 | 136.7415 | -81.4944 | 32260000 | 196.222 | -80.5156
32400000 | 136.0414 | -81.466 | 32400000 | 195.2794 | -80.459
32540000 | 135.2503 | -81.4004 | 32540000 | 194.243 | -80.4373
32680000 | 134.5803 | -81.3587 | 32680000 | 193.2252 | -80.3787
32820000 | 133.8928 | -81.2634 | 32820000 | 192.2393 | -80.3007
32960000 | 133.2172 | -81.2185 | 32960000 | 191.3204 | -80.2365
33100000 | 132.5621 | -81.145 | 33100000 | 190.3924 | -80.115
33240000 | 131.9106 | -81.0676 | 33240000 | 189.5383 | -80.0408
33380000 | 131.2665 | -80.9692 | 33380000 | 188.5708 | -79.9572
33520000 | 130.6468 | -80.8599 | 33520000 | 187.8095 | -79.8251
33660000 | 130.0629 | -80.7578 | 33660000 | 187.0315 | -79.7134
33800000 | 129.4563 | -80.6795 | 33800000 | 186.2855 | -79.5934
33940000 | 128.9115 | -80.5695 | 33940000 | 185.4782 | -79.4686
34080000 | 128.3731 | -80.4744 | 34080000 | 184.8327 | -79.3338
34220000 | 127.8694 | -80.3297 | 34220000 | 184.1672 | -79.2129
34360000 | 127.419 | -80.2134 | 34360000 | 183.4216 | -79.062
34500000 | 126.899 | -80.092 | 34500000 | 182.852 | -78.9173
34640000 | 126.4757 | -79.9795 | 34640000 | 182.3257 | -78.7608
34780000 | 126.096 | -79.8545 | 34780000 | 181.8925 | -78.6171
34920000 | 125.7096 | -79.7587 | 34920000 | 181.4103 | -78.4828
35060000 | 125.4113 | -79.6335 | 35060000 | 181.0759 | -78.3434
35200000 | 125.1073 | -79.533 | 35200000 | 180.9033 | -78.2187
35340000 | 124.8933 | -79.4253 | 35340000 | 180.6862 | -78.0976
35480000 | 124.6417 | -79.3761 | 35480000 | 180.4777 | -77.9916
35620000 | 124.4654 | -79.3123 | 35620000 | 180.397 | -77.9556
35760000 | 124.2261 | -79.3106 | 35760000 | 180.3116 | -77.916
35900000 | 124.0442 | -79.2817 | 35900000 | 180.2469 | -77.8868
36040000 | 123.8219 | -79.312 | 36040000 | 180.1876 | -77.9459
36180000 | 123.5897 | -79.3693 | 36180000 | 180.1066 | -77.9914
36320000 | 123.3032 | -79.4123 | 36320000 | 179.9121 | -78.0623
36460000 | 123.0644 | -79.5058 | 36460000 | 179.7181 | -78.1765
36600000 | 122.7073 | -79.5816 | 36600000 | 179.3807 | -78.2702
36740000 | 122.291 | -79.6876 | 36740000 | 178.9723 | -78.3956
36880000 | 121.9239 | -79.7734 | 36880000 | 178.5218 | -78.5556
37020000 | 121.4678 | -79.8819 | 37020000 | 177.9893 | -78.6749
37160000 | 120.8969 | -79.9795 | 37160000 | 177.3693 | -78.828
37300000 | 120.4324 | -80.0514 | 37300000 | 176.6012 | -78.9542
37440000 | 119.8534 | -80.1461 | 37440000 | 175.8784 | -79.0398
37580000 | 119.285 | -80.2019 | 37580000 | 175.1422 | -79.1566
37720000 | 118.6801 | -80.2527 | 37720000 | 174.313 | -79.2286
37860000 | 118.0989 | -80.2869 | 37860000 | 173.4731 | -79.3092
38000000 | 117.4708 | -80.3227 | 38000000 | 172.5669 | -79.3434
38140000 | 116.8824 | -80.3274 | 38140000 | 171.8037 | -79.3972
38280000 | 116.2731 | -80.3482 | 38280000 | 170.9094 | -79.4757
38420000 | 115.6693 | -80.3417 | 38420000 | 170.0389 | -79.4712
38560000 | 115.0584 | -80.3653 | 38560000 | 169.2037 | -79.505
38700000 | 114.4585 | -80.3371 | 38700000 | 168.3084 | -79.4885
38840000 | 113.8877 | -80.3441 | 38840000 | 167.4315 | -79.5371
38980000 | 113.2691 | -80.3112 | 38980000 | 166.5272 | -79.5184
39120000 | 112.6205 | -80.2966 | 39120000 | 165.6699 | -79.5072
39260000 | 112.0529 | -80.251 | 39260000 | 164.73 | -79.495
39400000 | 111.4263 | -80.1979 | 39400000 | 163.9168 | -79.4584
39540000 | 110.8538 | -80.1584 | 39540000 | 163.0104 | -79.4246
39680000 | 110.2452 | -80.0977 | 39680000 | 162.1502 | -79.373
39820000 | 109.6983 | -80.0319 | 39820000 | 161.2877 | -79.2886
39960000 | 109.1102 | -79.9623 | 39960000 | 160.4407 | -79.2229
40100000 | 108.5548 | -79.8699 | 40100000 | 159.6696 | -79.165
40240000 | 108.0109 | -79.8008 | 40240000 | 158.8306 | -79.0869
40380000 | 107.4985 | -79.7125 | 40380000 | 158.0541 | -78.9769
40520000 | 106.9494 | -79.61 | 40520000 | 157.2401 | -78.8834
40660000 | 106.4073 | -79.5138 | 40660000 | 156.4677 | -78.7749
40800000 | 105.9026 | -79.3977 | 40800000 | 155.7556 | -78.6623
40940000 | 105.3725 | -79.2951 | 40940000 | 155.0067 | -78.5416
41080000 | 104.8898 | -79.1705 | 41080000 | 154.2291 | -78.4338
41220000 | 104.421 | -79.051 | 41220000 | 153.4867 | -78.2934
41360000 | 103.9358 | -78.9228 | 41360000 | 152.8436 | -78.1459
41500000 | 103.4917 | -78.8034 | 41500000 | 152.1878 | -78.0104
41640000 | 103.0223 | -78.6596 | 41640000 | 151.5145 | -77.8724
41780000 | 102.5603 | -78.5238 | 41780000 | 150.8887 | -77.695
41920000 | 102.1688 | -78.3757 | 41920000 | 150.3203 | -77.5254
42060000 | 101.7715 | -78.2271 | 42060000 | 149.6909 | -77.3935
42200000 | 101.3323 | -78.0789 | 42200000 | 149.1472 | -77.2172
42340000 | 100.9397 | -77.9036 | 42340000 | 148.6436 | -77.0315
42480000 | 100.573 | -77.7774 | 42480000 | 148.1353 | -76.8468
42620000 | 100.2571 | -77.6253 | 42620000 | 147.6801 | -76.6812
42760000 | 99.8971 | -77.4626 | 42760000 | 147.1676 | -76.5003
42900000 | 99.5468 | -77.3118 | 42900000 | 146.7811 | -76.3316
43040000 | 99.2647 | -77.1362 | 43040000 | 146.3156 | -76.1652
43180000 | 98.9355 | -76.9908 | 43180000 | 145.9059 | -75.9778
43320000 | 98.6516 | -76.8392 | 43320000 | 145.5626 | -75.8207
43460000 | 98.382 | -76.6821 | 43460000 | 145.193 | -75.6319
43600000 | 98.0994 | -76.543 | 43600000 | 144.8311 | -75.4638
43740000 | 97.8377 | -76.4007 | 43740000 | 144.53 | -75.3017
43880000 | 97.5984 | -76.2544 | 43880000 | 144.2219 | -75.125
44020000 | 97.3618 | -76.1211 | 44020000 | 143.8867 | -74.9518
44160000 | 97.1098 | -75.9935 | 44160000 | 143.6776 | -74.7904
44300000 | 96.8704 | -75.8692 | 44300000 | 143.3265 | -74.6459
44440000 | 96.6397 | -75.7262 | 44440000 | 143.0961 | -74.4881
44580000 | 96.4297 | -75.5837 | 44580000 | 142.8482 | -74.3176
44720000 | 96.2157 | -75.471 | 44720000 | 142.5788 | -74.1999
44860000 | 96.0188 | -75.3571 | 44860000 | 142.4099 | -74.0628
45000000 | 95.8143 | -75.2482 | 45000000 | 142.1717 | -73.8863
45140000 | 95.5995 | -75.1132 | 45140000 | 141.9129 | -73.785
45280000 | 95.4051 | -75.0007 | 45280000 | 141.72 | -73.6685
45420000 | 95.2303 | -74.9101 | 45420000 | 141.4922 | -73.5337
45560000 | 95.0297 | -74.8003 | 45560000 | 141.2985 | -73.3879
45700000 | 94.8307 | -74.6867 | 45700000 | 141.1232 | -73.2922
45840000 | 94.6652 | -74.5952 | 45840000 | 140.9247 | -73.1707
45980000 | 94.4523 | -74.4962 | 45980000 | 140.6951 | -73.0562
46120000 | 94.2469 | -74.3891 | 46120000 | 140.4976 | -72.9411
46260000 | 94.0576 | -74.3153 | 46260000 | 140.3312 | -72.8532
46400000 | 93.8412 | -74.2216 | 46400000 | 140.09 | -72.7564
46540000 | 93.6565 | -74.1395 | 46540000 | 139.8293 | -72.6247
46680000 | 93.4258 | -74.0516 | 46680000 | 139.6242 | -72.5332
46820000 | 93.2421 | -73.974 | 46820000 | 139.3915 | -72.4328
46960000 | 93.0251 | -73.8697 | 46960000 | 139.1555 | -72.3417
47100000 | 92.789 | -73.7798 | 47100000 | 138.8118 | -72.2333
47240000 | 92.5654 | -73.6851 | 47240000 | 138.598 | -72.1277
47380000 | 92.3514 | -73.5823 | 47380000 | 138.336 | -72.0044
47520000 | 92.1292 | -73.465 | 47520000 | 138.0427 | -71.896
47660000 | 91.8785 | -73.3603 | 47660000 | 137.8185 | -71.7544
47800000 | 91.6736 | -73.2546 | 47800000 | 137.5213 | -71.6254
47940000 | 91.4422 | -73.1115 | 47940000 | 137.1895 | -71.4826
48080000 | 91.2211 | -72.9886 | 48080000 | 136.9288 | -71.3423
48220000 | 90.9952 | -72.8586 | 48220000 | 136.6606 | -71.1981
48360000 | 90.7957 | -72.715 | 48360000 | 136.3873 | -71.0314
48500000 | 90.5944 | -72.5724 | 48500000 | 136.0801 | -70.856
48640000 | 90.3898 | -72.4206 | 48640000 | 135.8692 | -70.7121
48780000 | 90.2129 | -72.2669 | 48780000 | 135.6248 | -70.5072
48920000 | 90.02 | -72.1085 | 48920000 | 135.3895 | -70.2933
49060000 | 89.8686 | -71.9416 | 49060000 | 135.1969 | -70.1019
49200000 | 89.7168 | -71.7542 | 49200000 | 134.9958 | -69.8999
49340000 | 89.5844 | -71.5683 | 49340000 | 134.8362 | -69.6863
49480000 | 89.4532 | -71.4036 | 49480000 | 134.6874 | -69.4439
49620000 | 89.3363 | -71.2203 | 49620000 | 134.5784 | -69.2299
49760000 | 89.264 | -71.0201 | 49760000 | 134.5186 | -68.9965
49900000 | 89.1805 | -70.8209 | 49900000 | 134.4162 | -68.7938
50040000 | 89.1133 | -70.6323 | 50040000 | 134.417 | -68.528
50180000 | 89.0851 | -70.4465 | 50180000 | 134.418 | -68.2804
50320000 | 89.0642 | -70.2481 | 50320000 | 134.5014 | -68.0573
50460000 | 89.0698 | -70.0549 | 50460000 | 134.5397 | -67.7713
50600000 | 89.0907 | -69.8676 | 50600000 | 134.6052 | -67.5368
50740000 | 89.1405 | -69.6952 | 50740000 | 134.8322 | -67.3083
50880000 | 89.2187 | -69.5152 | 50880000 | 135.0092 | -67.0531
51020000 | 89.2786 | -69.3356 | 51020000 | 135.2375 | -66.8264
51160000 | 89.3911 | -69.1691 | 51160000 | 135.498 | -66.5764
51300000 | 89.5002 | -69.0036 | 51300000 | 135.836 | -66.3743
51440000 | 89.6578 | -68.8539 | 51440000 | 136.1846 | -66.1446
51580000 | 89.8198 | -68.7424 | 51580000 | 136.5661 | -65.9547
51720000 | 89.9956 | -68.6061 | 51720000 | 136.9871 | -65.7541
51860000 | 90.1918 | -68.4907 | 51860000 | 137.4699 | -65.5589
52000000 | 90.3739 | -68.3878 | 52000000 | 137.9394 | -65.3751
52140000 | 90.5957 | -68.3115 | 52140000 | 138.438 | -65.2491
52280000 | 90.8133 | -68.248 | 52280000 | 139.0164 | -65.0801
52420000 | 91.036 | -68.1842 | 52420000 | 139.6116 | -64.9845
52560000 | 91.2689 | -68.1508 | 52560000 | 140.162 | -64.8577
52700000 | 91.5118 | -68.1246 | 52700000 | 140.7664 | -64.7618
52840000 | 91.7493 | -68.1116 | 52840000 | 141.3818 | -64.681
52980000 | 91.9508 | -68.1013 | 52980000 | 141.9621 | -64.6117
53120000 | 92.1715 | -68.1349 | 53120000 | 142.5891 | -64.5941
53260000 | 92.3903 | -68.1409 | 53260000 | 143.2042 | -64.5358
53400000 | 92.6048 | -68.1898 | 53400000 | 143.8303 | -64.5548
53540000 | 92.817 | -68.225 | 53540000 | 144.4714 | -64.524
53680000 | 93.0154 | -68.2828 | 53680000 | 145.035 | -64.5328
53820000 | 93.2097 | -68.3292 | 53820000 | 145.6895 | -64.5438
53960000 | 93.4236 | -68.4234 | 53960000 | 146.3053 | -64.5726
54100000 | 93.6025 | -68.4783 | 54100000 | 146.9346 | -64.5794
54240000 | 93.8076 | -68.5461 | 54240000 | 147.6175 | -64.6422
54380000 | 94.0001 | -68.6431 | 54380000 | 148.2985 | -64.7252
54520000 | 94.1842 | -68.7495 | 54520000 | 148.9718 | -64.7893
54660000 | 94.3714 | -68.8402 | 54660000 | 149.5775 | -64.8857
54800000 | 94.5506 | -68.9417 | 54800000 | 150.2803 | -64.9507
54940000 | 94.7457 | -69.0825 | 54940000 | 151.0001 | -65.0769
55080000 | 94.9168 | -69.2358 | 55080000 | 151.6318 | -65.2113
55220000 | 95.1207 | -69.3697 | 55220000 | 152.3133 | -65.3556
55360000 | 95.2636 | -69.5559 | 55360000 | 152.965 | -65.4986
55500000 | 95.4064 | -69.7222 | 55500000 | 153.6181 | -65.6979
55640000 | 95.5343 | -69.9005 | 55640000 | 154.2423 | -65.901
55780000 | 95.6616 | -70.0991 | 55780000 | 154.8435 | -66.0944
55920000 | 95.7775 | -70.3102 | 55920000 | 155.4068 | -66.339
56060000 | 95.8284 | -70.5119 | 56060000 | 155.9992 | -66.5968
56200000 | 95.8864 | -70.7424 | 56200000 | 156.4655 | -66.8591
56340000 | 95.9404 | -70.9486 | 56340000 | 156.8991 | -67.1131
56480000 | 95.9548 | -71.1945 | 56480000 | 157.3622 | -67.3904
56620000 | 95.9582 | -71.4267 | 56620000 | 157.7098 | -67.6666
56760000 | 95.9305 | -71.6529 | 56760000 | 158.0477 | -67.9093
56900000 | 95.8961 | -71.8676 | 56900000 | 158.4257 | -68.212
57040000 | 95.8098 | -72.1038 | 57040000 | 158.6684 | -68.5121
57180000 | 95.7561 | -72.3218 | 57180000 | 158.8369 | -68.8085
57320000 | 95.6645 | -72.5546 | 57320000 | 159.0579 | -69.0939
57460000 | 95.5735 | -72.7686 | 57460000 | 159.213 | -69.3816
57600000 | 95.4419 | -72.9987 | 57600000 | 159.3665 | -69.663
57740000 | 95.3302 | -73.2333 | 57740000 | 159.4305 | -69.9535
57880000 | 95.2278 | -73.442 | 57880000 | 159.578 | -70.2464
58020000 | 95.0535 | -73.6671 | 58020000 | 159.6377 | -70.5225
58160000 | 94.899 | -73.8939 | 58160000 | 159.6849 | -70.8299
58300000 | 94.7539 | -74.097 | 58300000 | 159.7161 | -71.122
58440000 | 94.598 | -74.3284 | 58440000 | 159.6734 | -71.4234
58580000 | 94.4221 | -74.5401 | 58580000 | 159.6616 | -71.7441
58720000 | 94.1834 | -74.7522 | 58720000 | 159.564 | -72.0368
58860000 | 93.9725 | -74.96 | 58860000 | 159.4508 | -72.3367
59000000 | 93.7691 | -75.1937 | 59000000 | 159.2981 | -72.6458
59140000 | 93.5027 | -75.3855 | 59140000 | 159.0866 | -72.9204
59280000 | 93.2499 | -75.5781 | 59280000 | 158.9029 | -73.2068
59420000 | 93.0013 | -75.7836 | 59420000 | 158.6663 | -73.4787
59560000 | 92.745 | -75.9782 | 59560000 | 158.3453 | -73.7568
59700000 | 92.4462 | -76.1548 | 59700000 | 158.0985 | -74.0207
59840000 | 92.1585 | -76.3466 | 59840000 | 157.7347 | -74.2849
59980000 | 91.9016 | -76.5065 | 59980000 | 157.5223 | -74.5367
60120000 | 91.5996 | -76.6741 | 60120000 | 157.1429 | -74.7723
60260000 | 91.3385 | -76.8364 | 60260000 | 156.8254 | -75.0113
60400000 | 91.0509 | -77.0029 | 60400000 | 156.4936 | -75.2385
60540000 | 90.7562 | -77.1377 | 60540000 | 156.1539 | -75.4766
60680000 | 90.4876 | -77.2851 | 60680000 | 155.8281 | -75.6968
60820000 | 90.1877 | -77.4243 | 60820000 | 155.4575 | -75.8725
60960000 | 89.8985 | -77.5909 | 60960000 | 155.1312 | -76.1121
61100000 | 89.632 | -77.7197 | 61100000 | 154.7608 | -76.332
61240000 | 89.3323 | -77.8679 | 61240000 | 154.3977 | -76.5547
61380000 | 89.0644 | -78.0056 | 61380000 | 154.0421 | -76.7575
61520000 | 88.7577 | -78.1374 | 61520000 | 153.5991 | -76.961
61660000 | 88.4737 | -78.2726 | 61660000 | 153.2127 | -77.1732
61800000 | 88.1386 | -78.3963 | 61800000 | 152.8207 | -77.3518
61940000 | 87.8459 | -78.5244 | 61940000 | 152.4196 | -77.548
62080000 | 87.5353 | -78.6376 | 62080000 | 151.994 | -77.7405
62220000 | 87.2411 | -78.7589 | 62220000 | 151.6124 | -77.9235
62360000 | 86.9244 | -78.8743 | 62360000 | 151.1422 | -78.1155
62500000 | 86.6325 | -78.9904 | 62500000 | 150.7286 | -78.2821
62640000 | 86.309 | -79.1027 | 62640000 | 150.2592 | -78.469
62780000 | 86.0226 | -79.209 | 62780000 | 149.8501 | -78.6016
62920000 | 85.6873 | -79.3051 | 62920000 | 149.3838 | -78.7861
63060000 | 85.3861 | -79.4112 | 63060000 | 148.8933 | -78.9309
63200000 | 85.0733 | -79.4943 | 63200000 | 148.4339 | -79.0709
63340000 | 84.7839 | -79.6038 | 63340000 | 148.0266 | -79.1973
63480000 | 84.4436 | -79.6619 | 63480000 | 147.5564 | -79.3461
63620000 | 84.1665 | -79.7427 | 63620000 | 147.1467 | -79.4759
63760000 | 83.8453 | -79.8241 | 63760000 | 146.7243 | -79.6018
63900000 | 83.584 | -79.898 | 63900000 | 146.2283 | -79.7153
64040000 | 83.2674 | -79.9546 | 64040000 | 145.8444 | -79.8502
64180000 | 82.9695 | -80.0327 | 64180000 | 145.3675 | -79.9515
64320000 | 82.6967 | -80.094 | 64320000 | 144.926 | -80.065
64460000 | 82.4109 | -80.18 | 64460000 | 144.5245 | -80.1516
64600000 | 82.1441 | -80.2203 | 64600000 | 144.1194 | -80.269
64740000 | 81.8626 | -80.2888 | 64740000 | 143.7004 | -80.3667
64880000 | 81.5866 | -80.3337 | 64880000 | 143.2276 | -80.4716
65020000 | 81.2925 | -80.3866 | 65020000 | 142.8638 | -80.5347
65160000 | 81.0208 | -80.4204 | 65160000 | 142.441 | -80.6185
65300000 | 80.7876 | -80.4753 | 65300000 | 142.0284 | -80.7192
65440000 | 80.5094 | -80.5044 | 65440000 | 141.6762 | -80.787
65580000 | 80.2379 | -80.5587 | 65580000 | 141.2672 | -80.8807
65720000 | 79.9867 | -80.603 | 65720000 | 140.9605 | -80.9306
65860000 | 79.7467 | -80.6329 | 65860000 | 140.5739 | -81.0102
66000000 | 79.5252 | -80.6691 | 66000000 | 140.237 | -81.0696
66140000 | 79.278 | -80.6985 | 66140000 | 139.8794 | -81.1352
66280000 | 79.061 | -80.731 | 66280000 | 139.6036 | -81.2046
66420000 | 78.8222 | -80.7677 | 66420000 | 139.296 | -81.2538
66560000 | 78.6032 | -80.8063 | 66560000 | 138.9991 | -81.3187
66700000 | 78.3855 | -80.8462 | 66700000 | 138.7398 | -81.3773
66840000 | 78.19 | -80.8823 | 66840000 | 138.4904 | -81.4323
66980000 | 77.9837 | -80.9232 | 66980000 | 138.2376 | -81.5046
67120000 | 77.801 | -80.9581 | 67120000 | 137.9681 | -81.5995
67260000 | 77.5826 | -81.0112 | 67260000 | 137.7281 | -81.6299
67400000 | 77.4097 | -81.0352 | 67400000 | 137.5159 | -81.6959
67540000 | 77.2232 | -81.0802 | 67540000 | 137.3298 | -81.7892
67680000 | 77.0507 | -81.1442 | 67680000 | 137.0811 | -81.8636
67820000 | 76.8611 | -81.1804 | 67820000 | 136.9352 | -81.935
67960000 | 76.7013 | -81.2334 | 67960000 | 136.6703 | -82.0052
68100000 | 76.5225 | -81.3087 | 68100000 | 136.578 | -82.1218
68240000 | 76.3451 | -81.3799 | 68240000 | 136.3558 | -82.2234
68380000 | 76.1649 | -81.4404 | 68380000 | 136.181 | -82.3134
68520000 | 76.0263 | -81.5247 | 68520000 | 135.9782 | -82.4278
68660000 | 75.8361 | -81.5961 | 68660000 | 135.8741 | -82.5442
68800000 | 75.6761 | -81.6821 | 68800000 | 135.6591 | -82.6569
68940000 | 75.4996 | -81.7569 | 68940000 | 135.4609 | -82.805
69080000 | 75.3075 | -81.8554 | 69080000 | 135.2565 | -82.9146
69220000 | 75.1272 | -81.9396 | 69220000 | 135.0455 | -83.0177
69360000 | 74.959 | -82.0461 | 69360000 | 134.8675 | -83.1846
69500000 | 74.7718 | -82.1544 | 69500000 | 134.652 | -83.3248
69640000 | 74.57 | -82.2292 | 69640000 | 134.4538 | -83.4626
69780000 | 74.3901 | -82.3576 | 69780000 | 134.183 | -83.6098
69920000 | 74.1788 | -82.4462 | 69920000 | 133.946 | -83.7743
70060000 | 73.9853 | -82.5395 | 70060000 | 133.6844 | -83.8996
70200000 | 73.7681 | -82.6571 | 70200000 | 133.4404 | -84.0533
70340000 | 73.5404 | -82.7756 | 70340000 | 133.1151 | -84.1998
70480000 | 73.3272 | -82.8723 | 70480000 | 132.7984 | -84.3532
70620000 | 73.1014 | -82.9886 | 70620000 | 132.4765 | -84.5135
70760000 | 72.8794 | -83.0751 | 70760000 | 132.2065 | -84.6581
70900000 | 72.6465 | -83.1822 | 70900000 | 131.8532 | -84.8193
71040000 | 72.4083 | -83.2926 | 71040000 | 131.5193 | -84.9704
71180000 | 72.1809 | -83.4025 | 71180000 | 131.1757 | -85.1067
71320000 | 71.9218 | -83.4977 | 71320000 | 130.8139 | -85.2656
71460000 | 71.6749 | -83.5841 | 71460000 | 130.4901 | -85.406
71600000 | 71.4358 | -83.6791 | 71600000 | 130.09 | -85.5604
71740000 | 71.1756 | -83.7851 | 71740000 | 129.7332 | -85.6916
71880000 | 70.9269 | -83.8655 | 71880000 | 129.3241 | -85.8103
72020000 | 70.6683 | -83.9843 | 72020000 | 128.9663 | -85.9544
72160000 | 70.4016 | -84.0688 | 72160000 | 128.5518 | -86.0976
72300000 | 70.1493 | -84.1485 | 72300000 | 128.1671 | -86.2064
72440000 | 69.8998 | -84.2212 | 72440000 | 127.7756 | -86.3317
72580000 | 69.6263 | -84.3005 | 72580000 | 127.3804 | -86.4425
72720000 | 69.3877 | -84.395 | 72720000 | 126.9502 | -86.5691
72860000 | 69.1112 | -84.4474 | 72860000 | 126.5819 | -86.6919
73000000 | 68.8544 | -84.5207 | 73000000 | 126.1389 | -86.7839
73140000 | 68.5852 | -84.5802 | 73140000 | 125.7187 | -86.8959
73280000 | 68.3423 | -84.6637 | 73280000 | 125.3185 | -86.9964
73420000 | 68.0741 | -84.7135 | 73420000 | 124.9328 | -87.0967
73560000 | 67.8238 | -84.778 | 73560000 | 124.5013 | -87.1943
73700000 | 67.556 | -84.8406 | 73700000 | 124.1045 | -87.2931
73840000 | 67.3055 | -84.9019 | 73840000 | 123.7005 | -87.3784
73980000 | 67.0358 | -84.9597 | 73980000 | 123.2952 | -87.4718
74120000 | 66.7864 | -85.0207 | 74120000 | 122.8749 | -87.5671
74260000 | 66.5227 | -85.0641 | 74260000 | 122.4726 | -87.6638
74400000 | 66.267 | -85.1201 | 74400000 | 122.0856 | -87.7325
74540000 | 66.0239 | -85.1794 | 74540000 | 121.6912 | -87.8011
74680000 | 65.7865 | -85.2189 | 74680000 | 121.2739 | -87.881
74820000 | 65.5183 | -85.2595 | 74820000 | 120.8678 | -87.9633
74960000 | 65.2586 | -85.3089 | 74960000 | 120.4874 | -88.0377
75100000 | 65.0247 | -85.3517 | 75100000 | 120.0785 | -88.1125
75240000 | 64.7651 | -85.3878 | 75240000 | 119.6672 | -88.1793
75380000 | 64.527 | -85.4376 | 75380000 | 119.2795 | -88.2252
75520000 | 64.2691 | -85.4763 | 75520000 | 118.9044 | -88.3072
75660000 | 64.0336 | -85.5081 | 75660000 | 118.5163 | -88.3826
75800000 | 63.7898 | -85.5436 | 75800000 | 118.1242 | -88.4428
75940000 | 63.5447 | -85.5732 | 75940000 | 117.7527 | -88.5015
76080000 | 63.3046 | -85.6058 | 76080000 | 117.3593 | -88.5494
76220000 | 63.0642 | -85.6398 | 76220000 | 116.994 | -88.6187
76360000 | 62.8427 | -85.6632 | 76360000 | 116.5979 | -88.6536
76500000 | 62.5921 | -85.6922 | 76500000 | 116.2519 | -88.7081
76640000 | 62.3554 | -85.7197 | 76640000 | 115.8631 | -88.7521
76780000 | 62.1316 | -85.7515 | 76780000 | 115.486 | -88.8097
76920000 | 61.8824 | -85.7644 | 76920000 | 115.1127 | -88.8508
77060000 | 61.6446 | -85.7905 | 77060000 | 114.7442 | -88.8972
77200000 | 61.435 | -85.8097 | 77200000 | 114.3783 | -88.9388
77340000 | 61.2073 | -85.8297 | 77340000 | 114.0063 | -88.9724
77480000 | 60.9612 | -85.8482 | 77480000 | 113.6554 | -88.9948
77620000 | 60.7411 | -85.8654 | 77620000 | 113.351 | -89.0577
77760000 | 60.5273 | -85.8873 | 77760000 | 112.9912 | -89.0963
77900000 | 60.2963 | -85.9047 | 77900000 | 112.6263 | -89.1231
78040000 | 60.0761 | -85.9076 | 78040000 | 112.2812 | -89.1519
78180000 | 59.855 | -85.9273 | 78180000 | 111.9454 | -89.1871
78320000 | 59.6363 | -85.9389 | 78320000 | 111.6234 | -89.2205
78460000 | 59.4074 | -85.9401 | 78460000 | 111.2813 | -89.2473
78600000 | 59.1959 | -85.9701 | 78600000 | 110.951 | -89.2798
78740000 | 58.9866 | -85.9638 | 78740000 | 110.6479 | -89.3041
78880000 | 58.7855 | -85.98 | 78880000 | 110.3005 | -89.3361
79020000 | 58.566 | -85.9786 | 79020000 | 109.9865 | -89.3682
79160000 | 58.3541 | -85.9909 | 79160000 | 109.6904 | -89.3865
79300000 | 58.1585 | -86.0045 | 79300000 | 109.3676 | -89.4303
79440000 | 57.9513 | -86.0224 | 79440000 | 109.0435 | -89.4502
79580000 | 57.7409 | -86.0142 | 79580000 | 108.7352 | -89.4681
79720000 | 57.5368 | -86.013 | 79720000 | 108.4218 | -89.4951
79860000 | 57.3278 | -86.0272 | 79860000 | 108.1142 | -89.5101
80000000 | 57.1234 | -86.0233 | 80000000 | 107.8266 | -89.5228

## Impedance (Re&Im)
Probe is fabricated with a 12 cm length of micro-coax attached to the probe head.
Impedance measurements have been taken with and without fixture compensation for this probe lead.
When measured with a 50resistor termination, the 12 cm micro-coax appears as a 15 nH series inductor.
Probe design
Lithium Nibate Composite
350 um pitch with 30 um kerfs
Piezo thickness 68 um
12 um paralene matching layer
E-Solder conductive epoxy backing
Un-compensated |  |  | Compensated
Frequency | Magnitude | Phase | Frequency | Magnitude | Phase
10000000 | 11.7517 | -436.7134 | 10000000 | 16.754 | -595.6246
10140000 | 11.8915 | -430.1553 | 10140000 | 16.7935 | -586.5862
10280000 | 11.8439 | -423.8316 | 10280000 | 17.1121 | -577.7484
10420000 | 12.0377 | -417.6348 | 10420000 | 17.2546 | -569.178
10560000 | 12.2564 | -411.6077 | 10560000 | 17.8367 | -560.7548
10700000 | 12.3744 | -405.6827 | 10700000 | 18.1329 | -552.5276
10840000 | 12.8313 | -399.8163 | 10840000 | 18.6563 | -544.4257
10980000 | 13.3747 | -394.341 | 10980000 | 19.6583 | -536.6956
11120000 | 13.7252 | -389.1973 | 11120000 | 20.4201 | -529.8567
11260000 | 14.1248 | -384.3592 | 11260000 | 20.8926 | -523.3949
11400000 | 14.0701 | -379.6683 | 11400000 | 20.8883 | -517.032
11540000 | 13.8768 | -374.9086 | 11540000 | 20.6986 | -510.693
11680000 | 13.7581 | -370.1688 | 11680000 | 20.3288 | -504.1381
11820000 | 13.633 | -365.291 | 11820000 | 20.1284 | -497.4062
11960000 | 13.6267 | -360.5179 | 11960000 | 20.0581 | -490.8376
12100000 | 13.6151 | -355.7133 | 12100000 | 20.2384 | -484.0015
12240000 | 13.792 | -351.0533 | 12240000 | 20.1742 | -477.621
12380000 | 13.9708 | -346.552 | 12380000 | 20.6303 | -471.3272
12520000 | 14.1585 | -342.0606 | 12520000 | 20.8112 | -465.201
12660000 | 14.3348 | -337.7535 | 12660000 | 21.2135 | -459.1588
12800000 | 14.5507 | -333.4949 | 12800000 | 21.5765 | -453.101
12940000 | 14.874 | -329.2769 | 12940000 | 21.9655 | -447.2679
13080000 | 15.1133 | -325.0943 | 13080000 | 22.4697 | -441.4388
13220000 | 15.4881 | -321.0698 | 13220000 | 22.9023 | -435.6638
13360000 | 15.7949 | -317.0687 | 13360000 | 23.6074 | -429.8354
13500000 | 16.1154 | -313.1027 | 13500000 | 24.2863 | -424.3848
13640000 | 16.5888 | -309.2554 | 13640000 | 25.0923 | -418.7514
13780000 | 17.0051 | -305.4357 | 13780000 | 25.6508 | -413.4517
13920000 | 17.5118 | -301.6824 | 13920000 | 26.6027 | -408.1441
14060000 | 17.9694 | -298.0193 | 14060000 | 27.4447 | -402.8132
14200000 | 18.5587 | -294.3089 | 14200000 | 28.5228 | -397.5685
14340000 | 19.2504 | -290.5861 | 14340000 | 29.6309 | -392.2186
14480000 | 19.9629 | -286.8856 | 14480000 | 30.8181 | -387.0782
14620000 | 20.8955 | -283.269 | 14620000 | 32.4994 | -381.7316
14760000 | 21.9766 | -279.7421 | 14760000 | 34.3994 | -376.6658
14900000 | 23.1944 | -276.3405 | 14900000 | 36.5282 | -371.8184
15040000 | 24.2907 | -272.9667 | 15040000 | 38.3545 | -366.9108
15180000 | 25.7441 | -269.7963 | 15180000 | 41.2747 | -362.2032
15320000 | 27.4279 | -266.8749 | 15320000 | 44.0813 | -357.9122
15460000 | 29.2276 | -263.9032 | 15460000 | 47.092 | -353.9423
15600000 | 31.3364 | -261.4664 | 15600000 | 50.6346 | -350.4792
15740000 | 33.6289 | -259.4575 | 15740000 | 54.8229 | -347.5476
15880000 | 35.7377 | -257.9818 | 15880000 | 58.5683 | -345.7382
16020000 | 37.5414 | -256.6971 | 16020000 | 61.7139 | -344.3112
16160000 | 39.2492 | -255.5485 | 16160000 | 64.8303 | -342.8576
16300000 | 41.1031 | -254.4682 | 16300000 | 68.4343 | -341.972
16440000 | 42.8857 | -254.0059 | 16440000 | 71.9868 | -341.5558
16580000 | 44.5494 | -253.8661 | 16580000 | 75.0428 | -341.7707
16720000 | 46.1126 | -254.1023 | 16720000 | 78.2885 | -343.2737
16860000 | 47.1084 | -255.0477 | 16860000 | 80.6806 | -345.5993
17000000 | 47.4035 | -256.2025 | 17000000 | 82.1911 | -348.542
17140000 | 47.2081 | -257.3936 | 17140000 | 82.0717 | -351.9718
17280000 | 46.0808 | -258.4515 | 17280000 | 80.7212 | -355.3945
17420000 | 44.5399 | -259.3208 | 17420000 | 78.3085 | -357.7465
17560000 | 42.2103 | -259.5177 | 17560000 | 74.8785 | -359.4609
17700000 | 39.8802 | -259.1716 | 17700000 | 70.7852 | -360.4179
17840000 | 37.9176 | -258.2936 | 17840000 | 66.9728 | -360.423
17980000 | 35.5631 | -257.0989 | 17980000 | 62.9243 | -359.4449
18120000 | 33.4645 | -255.4937 | 18120000 | 59.0532 | -357.9792
18260000 | 31.6642 | -253.6821 | 18260000 | 55.7259 | -356.0467
18400000 | 30.0246 | -251.9059 | 18400000 | 52.3411 | -353.7343
18540000 | 28.4035 | -249.9528 | 18540000 | 49.4417 | -351.1507
18680000 | 26.8645 | -247.6739 | 18680000 | 46.7442 | -348.3321
18820000 | 25.5096 | -245.5519 | 18820000 | 44.0901 | -344.961
18960000 | 24.2865 | -243.1089 | 18960000 | 41.3251 | -341.7852
19100000 | 23.2063 | -240.6352 | 19100000 | 39.2675 | -338.1887
19240000 | 22.3212 | -238.1918 | 19240000 | 37.5262 | -334.5325
19380000 | 21.6367 | -235.6223 | 19380000 | 36.2214 | -330.7336
19520000 | 21.0221 | -233.1061 | 19520000 | 35.1406 | -326.9781
19660000 | 20.8459 | -230.6256 | 19660000 | 34.5114 | -323.0029
19800000 | 20.4668 | -228.387 | 19800000 | 34.167 | -319.6116
19940000 | 20.4558 | -226.0257 | 19940000 | 33.8497 | -316.3962
20080000 | 20.3665 | -223.8636 | 20080000 | 33.5865 | -313.106
20220000 | 20.4626 | -221.6939 | 20220000 | 33.4483 | -309.8873
20360000 | 20.5244 | -219.5796 | 20360000 | 33.6305 | -306.889
20500000 | 20.6955 | -217.6223 | 20500000 | 34.0213 | -303.9759
20640000 | 21.0597 | -215.7022 | 20640000 | 34.5054 | -301.1332
20780000 | 21.386 | -213.9221 | 20780000 | 35.2011 | -298.5411
20920000 | 21.887 | -212.3123 | 20920000 | 36.0374 | -295.997
21060000 | 22.3777 | -210.8131 | 21060000 | 36.9384 | -294.2288
21200000 | 22.8199 | -209.455 | 21200000 | 37.8142 | -292.5673
21340000 | 23.1627 | -208.3598 | 21340000 | 38.4511 | -291.2291
21480000 | 23.4898 | -207.4438 | 21480000 | 38.8336 | -289.9734
21620000 | 23.6121 | -206.6236 | 21620000 | 39.4023 | -289.233
21760000 | 23.6521 | -205.8088 | 21760000 | 39.1492 | -288.4141
21900000 | 23.4276 | -205.0372 | 21900000 | 38.9241 | -287.7164
22040000 | 23.0141 | -204.3177 | 22040000 | 38.2533 | -287.1234
22180000 | 22.5399 | -203.3461 | 22180000 | 37.2903 | -286.1077
22320000 | 21.9936 | -202.4303 | 22320000 | 36.2706 | -284.9077
22460000 | 21.3875 | -201.187 | 22460000 | 35.0816 | -283.5159
22600000 | 20.6695 | -199.8468 | 22600000 | 34.0208 | -281.4429
22740000 | 20.1969 | -198.3708 | 22740000 | 33.0386 | -279.5953
22880000 | 19.808 | -196.9224 | 22880000 | 32.0613 | -277.3855
23020000 | 19.4099 | -195.4313 | 23020000 | 31.5767 | -275.3424
23160000 | 19.2115 | -194.0133 | 23160000 | 31.2203 | -273.3508
23300000 | 18.9376 | -192.5492 | 23300000 | 30.6854 | -271.4102
23440000 | 18.6696 | -191.1986 | 23440000 | 30.138 | -269.4882
23580000 | 18.5352 | -189.7653 | 23580000 | 29.6497 | -267.4092
23720000 | 18.2923 | -188.3471 | 23720000 | 29.402 | -265.0886
23860000 | 18.1711 | -186.9319 | 23860000 | 29.0247 | -263.2266
24000000 | 18.0847 | -185.4915 | 24000000 | 28.9476 | -261.037
24140000 | 17.9772 | -184.0947 | 24140000 | 28.7965 | -259.043
24280000 | 18.1065 | -182.6374 | 24280000 | 28.6834 | -257.0557
24420000 | 18.1372 | -181.2972 | 24420000 | 28.7457 | -255.1682
24560000 | 18.1208 | -180.0175 | 24560000 | 28.7946 | -253.3038
24700000 | 18.2285 | -178.7562 | 24700000 | 28.9137 | -251.483
24840000 | 18.2639 | -177.489 | 24840000 | 29.0125 | -249.8148
24980000 | 18.33 | -176.3153 | 24980000 | 29.1237 | -248.0237
25120000 | 18.4381 | -175.0992 | 25120000 | 29.0699 | -246.2137
25260000 | 18.4209 | -173.8623 | 25260000 | 29.3926 | -244.5239
25400000 | 18.6117 | -172.6912 | 25400000 | 29.5759 | -242.7391
25540000 | 18.8084 | -171.4962 | 25540000 | 29.8684 | -241.1017
25680000 | 19.0178 | -170.4022 | 25680000 | 30.3116 | -239.6137
25820000 | 19.2328 | -169.339 | 25820000 | 30.622 | -238.1708
25960000 | 19.4598 | -168.4142 | 25960000 | 31.1046 | -236.8643
26100000 | 19.6373 | -167.5338 | 26100000 | 31.3288 | -235.5849
26240000 | 19.6923 | -166.6847 | 26240000 | 31.5502 | -234.5726
26380000 | 19.7139 | -165.8203 | 26380000 | 31.6305 | -233.3047
26520000 | 19.6985 | -164.8869 | 26520000 | 31.3998 | -232.0443
26660000 | 19.7048 | -163.9111 | 26660000 | 31.4272 | -230.7347
26800000 | 19.7889 | -162.9012 | 26800000 | 31.6042 | -229.404
26940000 | 19.9353 | -161.8656 | 26940000 | 31.8129 | -227.9897
27080000 | 20.0553 | -160.9148 | 27080000 | 32.24 | -226.5944
27220000 | 20.2317 | -160.0175 | 27220000 | 32.4741 | -225.4671
27360000 | 20.4913 | -159.1956 | 27360000 | 32.8916 | -224.3572
27500000 | 20.6326 | -158.4126 | 27500000 | 33.2118 | -223.3403
27640000 | 20.7811 | -157.7084 | 27640000 | 33.482 | -222.2227
27780000 | 20.9277 | -156.9003 | 27780000 | 33.7705 | -221.3823
27920000 | 21.062 | -156.2054 | 27920000 | 33.926 | -220.5723
28060000 | 21.1623 | -155.4926 | 28060000 | 34.0819 | -219.7701
28200000 | 21.2761 | -154.8709 | 28200000 | 34.341 | -218.9354
28340000 | 21.2667 | -154.2349 | 28340000 | 34.349 | -218.1631
28480000 | 21.2686 | -153.6037 | 28480000 | 34.2825 | -217.3882
28620000 | 21.1902 | -152.9322 | 28620000 | 34.2824 | -216.5849
28760000 | 21.2128 | -152.2798 | 28760000 | 34.2952 | -215.7844
28900000 | 21.1057 | -151.6079 | 28900000 | 34.1327 | -214.9694
29040000 | 21.0859 | -150.8935 | 29040000 | 34.0314 | -214.093
29180000 | 21.0007 | -150.1791 | 29180000 | 33.8748 | -213.2057
29320000 | 20.884 | -149.5031 | 29320000 | 33.7507 | -212.1818
29460000 | 20.8371 | -148.746 | 29460000 | 33.5309 | -211.3154
29600000 | 20.7817 | -148.0209 | 29600000 | 33.4291 | -210.3902
29740000 | 20.6806 | -147.2834 | 29740000 | 33.269 | -209.3428
29880000 | 20.6565 | -146.5359 | 29880000 | 33.2406 | -208.3363
30020000 | 20.6372 | -145.6995 | 30020000 | 33.1664 | -207.264
30160000 | 20.6709 | -144.9535 | 30160000 | 33.2454 | -206.2351
30300000 | 20.6556 | -144.227 | 30300000 | 33.3531 | -205.1961
30440000 | 20.7382 | -143.5187 | 30440000 | 33.4656 | -204.2904
30580000 | 20.8398 | -142.8468 | 30580000 | 33.6523 | -203.4755
30720000 | 20.8796 | -142.2192 | 30720000 | 33.7141 | -202.5999
30860000 | 20.9455 | -141.6238 | 30860000 | 33.7389 | -201.8846
31000000 | 20.9317 | -141.0656 | 31000000 | 33.7656 | -201.177
31140000 | 20.9104 | -140.4738 | 31140000 | 33.6736 | -200.4333
31280000 | 20.8595 | -139.9069 | 31280000 | 33.5449 | -199.7672
31420000 | 20.7786 | -139.2803 | 31420000 | 33.4493 | -199.0335
31560000 | 20.6611 | -138.6937 | 31560000 | 33.2011 | -198.1558
31700000 | 20.5842 | -138.0221 | 31700000 | 32.9696 | -197.3041
31840000 | 20.4653 | -137.3528 | 31840000 | 32.8135 | -196.3916
31980000 | 20.3516 | -136.6775 | 31980000 | 32.6328 | -195.5243
32120000 | 20.2796 | -135.9471 | 32120000 | 32.4782 | -194.5142
32260000 | 20.2849 | -135.2314 | 32260000 | 32.4154 | -193.5106
32400000 | 20.1912 | -134.5327 | 32400000 | 32.3557 | -192.4532
32540000 | 20.2307 | -133.7527 | 32540000 | 32.3333 | -191.514
32680000 | 20.2514 | -133.1002 | 32680000 | 32.3887 | -190.4977
32820000 | 20.2857 | -132.35 | 32820000 | 32.3258 | -189.5147
32960000 | 20.3116 | -131.6625 | 32960000 | 32.6099 | -188.5662
33100000 | 20.4206 | -130.9872 | 33100000 | 32.6822 | -187.5698
33240000 | 20.5643 | -130.2753 | 33240000 | 32.7768 | -186.6171
33380000 | 20.6069 | -129.644 | 33380000 | 32.9075 | -185.7306
33520000 | 20.7596 | -128.9978 | 33520000 | 33.1593 | -184.8581
33660000 | 20.8376 | -128.4036 | 33660000 | 33.4242 | -184.0105
33800000 | 21.0225 | -127.7756 | 33800000 | 33.7336 | -183.11
33940000 | 21.1428 | -127.2217 | 33940000 | 33.8934 | -182.3773
34080000 | 21.3129 | -126.6544 | 34080000 | 34.1766 | -181.5558
34220000 | 21.4596 | -126.0745 | 34220000 | 34.4715 | -180.8518
34360000 | 21.6543 | -125.5604 | 34360000 | 34.8123 | -180.1181
34500000 | 21.8192 | -125.0038 | 34500000 | 35.1148 | -179.4556
34640000 | 21.9793 | -124.5668 | 34640000 | 35.4779 | -178.8393
34780000 | 22.2312 | -124.1089 | 34780000 | 35.8259 | -178.2641
34920000 | 22.3767 | -123.7016 | 34920000 | 36.2387 | -177.7809
35060000 | 22.5514 | -123.3291 | 35060000 | 36.6202 | -177.3932
35200000 | 22.7387 | -123.0264 | 35200000 | 37.0385 | -177.0191
35340000 | 22.878 | -122.7327 | 35340000 | 37.2769 | -176.7416
35480000 | 22.9678 | -122.4786 | 35480000 | 37.5323 | -176.605
35620000 | 23.0587 | -122.2854 | 35620000 | 37.6247 | -176.4922
35760000 | 23.0766 | -122.0796 | 35760000 | 37.7695 | -176.3452
35900000 | 23.0481 | -121.8792 | 35900000 | 37.7216 | -176.2818
36040000 | 22.9448 | -121.6904 | 36040000 | 37.7404 | -176.2462
36180000 | 22.7977 | -121.4597 | 36180000 | 37.4563 | -176.1125
36320000 | 22.6458 | -121.201 | 36320000 | 37.2398 | -175.9756
36460000 | 22.47 | -120.9853 | 36460000 | 36.8531 | -175.7505
36600000 | 22.2344 | -120.6488 | 36600000 | 36.4594 | -175.5562
36740000 | 21.9427 | -120.3413 | 36740000 | 35.9383 | -175.3387
36880000 | 21.7009 | -119.952 | 36880000 | 35.4678 | -174.8912
37020000 | 21.383 | -119.5759 | 37020000 | 34.9435 | -174.4917
37160000 | 21.0842 | -119.1179 | 37160000 | 34.4106 | -173.9509
37300000 | 20.8421 | -118.6088 | 37300000 | 33.8887 | -173.3948
37440000 | 20.5632 | -118.0732 | 37440000 | 33.3831 | -172.7139
37580000 | 20.3399 | -117.538 | 37580000 | 32.9513 | -171.9789
37720000 | 20.1373 | -116.9689 | 37720000 | 32.6064 | -171.2892
37860000 | 19.9828 | -116.4103 | 37860000 | 32.1032 | -170.4838
38000000 | 19.7937 | -115.8235 | 38000000 | 31.8952 | -169.6631
38140000 | 19.6505 | -115.2288 | 38140000 | 31.6019 | -168.8953
38280000 | 19.4828 | -114.6651 | 38280000 | 31.262 | -168.0802
38420000 | 19.3818 | -114.0387 | 38420000 | 31.0715 | -167.1754
38560000 | 19.2794 | -113.479 | 38560000 | 30.8132 | -166.3971
38700000 | 19.1997 | -112.8764 | 38700000 | 30.6443 | -165.4903
38840000 | 19.1342 | -112.2575 | 38840000 | 30.4464 | -164.6334
38980000 | 19.0567 | -111.6212 | 38980000 | 30.2394 | -163.7327
39120000 | 19.044 | -111.0298 | 39120000 | 30.1747 | -162.852
39260000 | 18.9744 | -110.4709 | 39260000 | 30.0321 | -161.9858
39400000 | 18.9362 | -109.8302 | 39400000 | 30.057 | -161.1717
39540000 | 18.9582 | -109.2287 | 39540000 | 29.9344 | -160.2876
39680000 | 18.9693 | -108.6248 | 39680000 | 29.9475 | -159.3759
39820000 | 19.0197 | -108.0403 | 39820000 | 29.9097 | -158.538
39960000 | 19.0537 | -107.4846 | 39960000 | 30.0072 | -157.6786
40100000 | 19.0895 | -106.8898 | 40100000 | 30.0763 | -156.8466
40240000 | 19.1394 | -106.3308 | 40240000 | 30.1095 | -155.936
40380000 | 19.2334 | -105.7681 | 40380000 | 30.1999 | -155.1094
40520000 | 19.2903 | -105.1988 | 40520000 | 30.3229 | -154.2977
40660000 | 19.3804 | -104.6457 | 40660000 | 30.4681 | -153.4487
40800000 | 19.4957 | -104.1021 | 40800000 | 30.5972 | -152.6856
40940000 | 19.6027 | -103.5877 | 40940000 | 30.7567 | -151.8528
41080000 | 19.6853 | -103.0558 | 41080000 | 30.9766 | -151.1166
41220000 | 19.8512 | -102.5157 | 41220000 | 31.2101 | -150.3744
41360000 | 19.9586 | -102.033 | 41360000 | 31.3655 | -149.6408
41500000 | 20.1148 | -101.5345 | 41500000 | 31.5961 | -148.8798
41640000 | 20.2666 | -100.999 | 41640000 | 31.8397 | -148.14
41780000 | 20.4432 | -100.5599 | 41780000 | 32.1677 | -147.4481
41920000 | 20.5928 | -100.0758 | 41920000 | 32.4329 | -146.7503
42060000 | 20.7702 | -99.6066 | 42060000 | 32.7339 | -146.1207
42200000 | 20.914 | -99.1941 | 42200000 | 32.99 | -145.499
42340000 | 21.1303 | -98.7575 | 42340000 | 33.3887 | -144.8676
42480000 | 21.3169 | -98.3306 | 42480000 | 33.629 | -144.2668
42620000 | 21.5176 | -97.9017 | 42620000 | 34.0148 | -143.682
42760000 | 21.675 | -97.5562 | 42760000 | 34.3435 | -143.1499
42900000 | 21.8721 | -97.1581 | 42900000 | 34.6871 | -142.5507
43040000 | 22.0626 | -96.808 | 43040000 | 34.9753 | -142.0929
43180000 | 22.2442 | -96.4473 | 43180000 | 35.3238 | -141.5768
43320000 | 22.4791 | -96.0704 | 43320000 | 35.717 | -141.1154
43460000 | 22.662 | -95.7664 | 43460000 | 35.9969 | -140.6228
43600000 | 22.8321 | -95.418 | 43600000 | 36.3571 | -140.1929
43740000 | 23.0031 | -95.1284 | 43740000 | 36.6836 | -139.8109
43880000 | 23.1508 | -94.8186 | 43880000 | 37.0168 | -139.4134
44020000 | 23.3576 | -94.5187 | 44020000 | 37.3687 | -138.9875
44160000 | 23.521 | -94.2564 | 44160000 | 37.6309 | -138.604
44300000 | 23.6527 | -93.9511 | 44300000 | 37.9654 | -138.2288
44440000 | 23.8486 | -93.6876 | 44440000 | 38.2766 | -137.9109
44580000 | 23.9967 | -93.4265 | 44580000 | 38.5549 | -137.5227
44720000 | 24.1466 | -93.1727 | 44720000 | 38.8357 | -137.2211
44860000 | 24.2828 | -92.8965 | 44860000 | 39.0973 | -136.8899
45000000 | 24.4256 | -92.6365 | 45000000 | 39.3663 | -136.5803
45140000 | 24.5495 | -92.4168 | 45140000 | 39.689 | -136.3074
45280000 | 24.7015 | -92.1598 | 45280000 | 39.9361 | -135.9481
45420000 | 24.8009 | -91.9495 | 45420000 | 40.1371 | -135.7092
45560000 | 24.9349 | -91.6945 | 45560000 | 40.3721 | -135.4357
45700000 | 25.079 | -91.4701 | 45700000 | 40.6048 | -135.1703
45840000 | 25.1503 | -91.2323 | 45840000 | 40.8404 | -134.8275
45980000 | 25.271 | -91.0335 | 45980000 | 40.9938 | -134.597
46120000 | 25.3484 | -90.7978 | 46120000 | 41.2223 | -134.3098
46260000 | 25.4609 | -90.5507 | 46260000 | 41.3647 | -134.0327
46400000 | 25.5006 | -90.3395 | 46400000 | 41.5593 | -133.7338
46540000 | 25.6051 | -90.1056 | 46540000 | 41.7637 | -133.4969
46680000 | 25.6922 | -89.8629 | 46680000 | 41.8567 | -133.144
46820000 | 25.7586 | -89.62 | 46820000 | 42.0629 | -132.8866
46960000 | 25.8558 | -89.3493 | 46960000 | 42.2346 | -132.536
47100000 | 25.9486 | -89.1266 | 47100000 | 42.446 | -132.2687
47240000 | 26.0465 | -88.8442 | 47240000 | 42.5381 | -131.9207
47380000 | 26.1156 | -88.6162 | 47380000 | 42.7371 | -131.6049
47520000 | 26.2254 | -88.3092 | 47520000 | 42.9033 | -131.2445
47660000 | 26.3282 | -88.0594 | 47660000 | 43.1277 | -130.8287
47800000 | 26.4287 | -87.7973 | 47800000 | 43.3196 | -130.5272
47940000 | 26.5329 | -87.5307 | 47940000 | 43.4977 | -130.1385
48080000 | 26.6885 | -87.2682 | 48080000 | 43.7567 | -129.7338
48220000 | 26.8105 | -86.9795 | 48220000 | 44.0537 | -129.3447
48360000 | 26.9512 | -86.7281 | 48360000 | 44.356 | -128.9957
48500000 | 27.1227 | -86.4338 | 48500000 | 44.6676 | -128.6147
48640000 | 27.2749 | -86.1912 | 48640000 | 44.8918 | -128.2212
48780000 | 27.4626 | -85.9364 | 48780000 | 45.2311 | -127.8278
48920000 | 27.6723 | -85.6865 | 48920000 | 45.6407 | -127.4435
49060000 | 27.8821 | -85.4301 | 49060000 | 46.0317 | -127.0864
49200000 | 28.0746 | -85.2312 | 49200000 | 46.3729 | -126.766
49340000 | 28.3051 | -85.008 | 49340000 | 46.812 | -126.4357
49480000 | 28.5363 | -84.8053 | 49480000 | 47.279 | -126.1296
49620000 | 28.7729 | -84.6094 | 49620000 | 47.7099 | -125.872
49760000 | 29.0052 | -84.4341 | 49760000 | 48.1947 | -125.5405
49900000 | 29.2507 | -84.2588 | 49900000 | 48.696 | -125.3567
50040000 | 29.5375 | -84.1167 | 50040000 | 49.1731 | -125.0579
50180000 | 29.7998 | -83.9815 | 50180000 | 49.6965 | -124.8577
50320000 | 30.0524 | -83.8433 | 50320000 | 50.2632 | -124.706
50460000 | 30.3671 | -83.7568 | 50460000 | 50.8623 | -124.5658
50600000 | 30.6601 | -83.6727 | 50600000 | 51.4798 | -124.4447
50740000 | 30.953 | -83.6084 | 50740000 | 52.0427 | -124.3889
50880000 | 31.206 | -83.5727 | 50880000 | 52.638 | -124.2997
51020000 | 31.4846 | -83.5721 | 51020000 | 53.2533 | -124.3239
51160000 | 31.7864 | -83.5542 | 51160000 | 53.8488 | -124.343
51300000 | 32.0607 | -83.5887 | 51300000 | 54.4912 | -124.4529
51440000 | 32.3326 | -83.6272 | 51440000 | 55.0884 | -124.5539
51580000 | 32.5863 | -83.7224 | 51580000 | 55.6966 | -124.7031
51720000 | 32.8217 | -83.8029 | 51720000 | 56.2798 | -124.8776
51860000 | 33.0519 | -83.9181 | 51860000 | 56.9262 | -125.099
52000000 | 33.283 | -84.0616 | 52000000 | 57.3959 | -125.4051
52140000 | 33.5022 | -84.1899 | 52140000 | 57.9929 | -125.7376
52280000 | 33.6452 | -84.3604 | 52280000 | 58.5671 | -126.0923
52420000 | 33.7945 | -84.5489 | 52420000 | 59.0761 | -126.4595
52560000 | 33.9574 | -84.7262 | 52560000 | 59.4948 | -126.9413
52700000 | 34.0831 | -84.9219 | 52700000 | 59.9889 | -127.3542
52840000 | 34.2008 | -85.1101 | 52840000 | 60.4309 | -127.7888
52980000 | 34.2624 | -85.3412 | 52980000 | 60.8151 | -128.3411
53120000 | 34.3128 | -85.565 | 53120000 | 61.186 | -128.8174
53260000 | 34.3658 | -85.7675 | 53260000 | 61.5172 | -129.3076
53400000 | 34.4098 | -85.9874 | 53400000 | 61.8796 | -129.8866
53540000 | 34.4314 | -86.1997 | 53540000 | 62.1211 | -130.4262
53680000 | 34.4393 | -86.4118 | 53680000 | 62.3546 | -131.0154
53820000 | 34.3957 | -86.6395 | 53820000 | 62.607 | -131.5916
53960000 | 34.3797 | -86.8261 | 53960000 | 62.8602 | -132.1291
54100000 | 34.3386 | -87.0582 | 54100000 | 63.043 | -132.7592
54240000 | 34.299 | -87.3 | 54240000 | 63.1462 | -133.3676
54380000 | 34.249 | -87.5202 | 54380000 | 63.2895 | -134.0763
54520000 | 34.1689 | -87.773 | 54520000 | 63.4881 | -134.7261
54660000 | 34.0481 | -88.0182 | 54660000 | 63.5425 | -135.3569
54800000 | 33.9644 | -88.265 | 54800000 | 63.5611 | -136.169
54940000 | 33.8141 | -88.5075 | 54940000 | 63.6004 | -136.9206
55080000 | 33.6703 | -88.7304 | 55080000 | 63.5444 | -137.6887
55220000 | 33.507 | -88.9966 | 55220000 | 63.557 | -138.472
55360000 | 33.32 | -89.2109 | 55360000 | 63.3547 | -139.2459
55500000 | 33.0731 | -89.4851 | 55500000 | 63.2066 | -139.9959
55640000 | 32.8522 | -89.7322 | 55640000 | 62.9843 | -140.8061
55780000 | 32.5653 | -89.9211 | 55780000 | 62.6785 | -141.5925
55920000 | 32.2556 | -90.1525 | 55920000 | 62.3173 | -142.3779
56060000 | 31.943 | -90.3809 | 56060000 | 61.9398 | -143.1124
56200000 | 31.623 | -90.5393 | 56200000 | 61.5659 | -143.8983
56340000 | 31.3123 | -90.6828 | 56340000 | 61.0434 | -144.515
56480000 | 30.9361 | -90.8074 | 56480000 | 60.5003 | -145.2911
56620000 | 30.5921 | -90.9488 | 56620000 | 60.0048 | -145.8651
56760000 | 30.2235 | -91.0669 | 56760000 | 59.3528 | -146.4929
56900000 | 29.8491 | -91.1348 | 56900000 | 58.8033 | -147.0266
57040000 | 29.4643 | -91.187 | 57040000 | 58.1273 | -147.6333
57180000 | 29.0608 | -91.2491 | 57180000 | 57.5183 | -148.0829
57320000 | 28.6967 | -91.2686 | 57320000 | 56.7371 | -148.5901
57460000 | 28.2938 | -91.2975 | 57460000 | 56.1111 | -149.0186
57600000 | 27.9017 | -91.2972 | 57600000 | 55.373 | -149.4178
57740000 | 27.5211 | -91.2766 | 57740000 | 54.713 | -149.8178
57880000 | 27.1475 | -91.2722 | 57880000 | 53.8528 | -150.2324
58020000 | 26.7386 | -91.2268 | 58020000 | 53.2322 | -150.4836
58160000 | 26.3356 | -91.1773 | 58160000 | 52.4405 | -150.862
58300000 | 25.979 | -91.1263 | 58300000 | 51.6684 | -151.1325
58440000 | 25.5691 | -91.0793 | 58440000 | 50.8136 | -151.4038
58580000 | 25.1491 | -90.9915 | 58580000 | 50.0072 | -151.5796
58720000 | 24.7591 | -90.8715 | 58720000 | 49.2093 | -151.7542
58860000 | 24.3691 | -90.775 | 58860000 | 48.4082 | -151.9275
59000000 | 23.9863 | -90.6266 | 59000000 | 47.5708 | -151.9763
59140000 | 23.589 | -90.5072 | 59140000 | 46.7336 | -152.166
59280000 | 23.2039 | -90.3332 | 59280000 | 45.9095 | -152.1324
59420000 | 22.8261 | -90.1874 | 59420000 | 45.1504 | -152.0864
59560000 | 22.4829 | -89.9631 | 59560000 | 44.2719 | -152.0705
59700000 | 22.116 | -89.8008 | 59700000 | 43.5068 | -151.9702
59840000 | 21.7827 | -89.5714 | 59840000 | 42.7272 | -151.8636
59980000 | 21.4515 | -89.3787 | 59980000 | 41.9828 | -151.8419
60120000 | 21.1227 | -89.1698 | 60120000 | 41.2036 | -151.6432
60260000 | 20.7978 | -88.9471 | 60260000 | 40.5754 | -151.456
60400000 | 20.5004 | -88.7155 | 60400000 | 39.9233 | -151.3523
60540000 | 20.2101 | -88.5075 | 60540000 | 39.1878 | -151.1272
60680000 | 19.9242 | -88.2756 | 60680000 | 38.4628 | -150.9445
60820000 | 19.6109 | -88.0513 | 60820000 | 37.881 | -150.7933
60960000 | 19.3422 | -87.8291 | 60960000 | 37.2469 | -150.6155
61100000 | 19.0708 | -87.6076 | 61100000 | 36.5161 | -150.3888
61240000 | 18.7928 | -87.3788 | 61240000 | 35.9534 | -150.161
61380000 | 18.5089 | -87.1191 | 61380000 | 35.3087 | -149.9233
61520000 | 18.2239 | -86.8702 | 61520000 | 34.6333 | -149.7024
61660000 | 17.9796 | -86.6144 | 61660000 | 34.0549 | -149.4038
61800000 | 17.7095 | -86.3688 | 61800000 | 33.4299 | -149.0858
61940000 | 17.4846 | -86.0986 | 61940000 | 32.8655 | -148.8668
62080000 | 17.2263 | -85.8469 | 62080000 | 32.2261 | -148.5434
62220000 | 16.9668 | -85.5793 | 62220000 | 31.7465 | -148.2321
62360000 | 16.7323 | -85.3197 | 62360000 | 31.1562 | -147.8926
62500000 | 16.545 | -85.0532 | 62500000 | 30.6414 | -147.5959
62640000 | 16.3144 | -84.7795 | 62640000 | 30.0954 | -147.185
62780000 | 16.0728 | -84.4932 | 62780000 | 29.6014 | -146.8997
62920000 | 15.8876 | -84.2135 | 62920000 | 29.0954 | -146.5021
63060000 | 15.6701 | -83.9321 | 63060000 | 28.6238 | -146.149
63200000 | 15.506 | -83.6798 | 63200000 | 28.1095 | -145.7618
63340000 | 15.3125 | -83.3771 | 63340000 | 27.737 | -145.4036
63480000 | 15.1338 | -83.1071 | 63480000 | 27.2792 | -145.0482
63620000 | 14.9663 | -82.8337 | 63620000 | 26.8387 | -144.6678
63760000 | 14.816 | -82.5711 | 63760000 | 26.47 | -144.2379
63900000 | 14.6577 | -82.3288 | 63900000 | 26.0725 | -143.919
64040000 | 14.4969 | -82.0288 | 64040000 | 25.7363 | -143.4924
64180000 | 14.3488 | -81.7535 | 64180000 | 25.3571 | -143.1446
64320000 | 14.2269 | -81.4733 | 64320000 | 25.0367 | -142.7549
64460000 | 14.0699 | -81.2201 | 64460000 | 24.6901 | -142.4144
64600000 | 13.9328 | -80.9381 | 64600000 | 24.3181 | -142.0057
64740000 | 13.8098 | -80.6835 | 64740000 | 24.0536 | -141.6301
64880000 | 13.6875 | -80.4236 | 64880000 | 23.7112 | -141.2518
65020000 | 13.5765 | -80.181 | 65020000 | 23.4864 | -140.8766
65160000 | 13.4591 | -79.9057 | 65160000 | 23.1998 | -140.5424
65300000 | 13.3626 | -79.6584 | 65300000 | 22.9022 | -140.183
65440000 | 13.2502 | -79.4033 | 65440000 | 22.7377 | -139.8047
65580000 | 13.1512 | -79.1622 | 65580000 | 22.4603 | -139.4858
65720000 | 13.0398 | -78.9115 | 65720000 | 22.1989 | -139.1759
65860000 | 12.96 | -78.6889 | 65860000 | 21.9644 | -138.8462
66000000 | 12.8898 | -78.4778 | 66000000 | 21.7585 | -138.5534
66140000 | 12.7817 | -78.243 | 66140000 | 21.55 | -138.2095
66280000 | 12.7242 | -78.0216 | 66280000 | 21.3548 | -137.9296
66420000 | 12.6348 | -77.7945 | 66420000 | 21.1863 | -137.6278
66560000 | 12.5615 | -77.5828 | 66560000 | 21.028 | -137.3908
66700000 | 12.4779 | -77.3774 | 66700000 | 20.8548 | -137.195
66840000 | 12.3951 | -77.2085 | 66840000 | 20.5749 | -136.9344
66980000 | 12.3005 | -77.0115 | 66980000 | 20.3992 | -136.7438
67120000 | 12.2222 | -76.8139 | 67120000 | 20.1939 | -136.4828
67260000 | 12.1433 | -76.6425 | 67260000 | 20.0198 | -136.2556
67400000 | 12.036 | -76.4687 | 67400000 | 19.8155 | -136.0414
67540000 | 11.9707 | -76.2901 | 67540000 | 19.6209 | -135.9136
67680000 | 11.8777 | -76.1331 | 67680000 | 19.3948 | -135.7442
67820000 | 11.7664 | -75.9619 | 67820000 | 19.1895 | -135.5569
67960000 | 11.6738 | -75.7999 | 67960000 | 18.9474 | -135.3875
68100000 | 11.5442 | -75.6482 | 68100000 | 18.68 | -135.2702
68240000 | 11.4392 | -75.4849 | 68240000 | 18.4763 | -135.1104
68380000 | 11.3492 | -75.3288 | 68380000 | 18.2139 | -134.9579
68520000 | 11.2274 | -75.1881 | 68520000 | 17.9101 | -134.8015
68660000 | 11.0786 | -75.0122 | 68660000 | 17.6134 | -134.6883
68800000 | 10.9697 | -74.8691 | 68800000 | 17.3316 | -134.4764
68940000 | 10.8284 | -74.7022 | 68940000 | 17.0099 | -134.4158
69080000 | 10.683 | -74.5391 | 69080000 | 16.6501 | -134.2036
69220000 | 10.5295 | -74.3986 | 69220000 | 16.3428 | -134.0482
69360000 | 10.381 | -74.216 | 69360000 | 16.0104 | -133.9362
69500000 | 10.2456 | -74.0489 | 69500000 | 15.6616 | -133.6891
69640000 | 10.0805 | -73.8838 | 69640000 | 15.2567 | -133.5543
69780000 | 9.9226 | -73.7068 | 69780000 | 14.9187 | -133.3432
69920000 | 9.7732 | -73.5235 | 69920000 | 14.5699 | -133.1379
70060000 | 9.5945 | -73.3332 | 70060000 | 14.1609 | -132.9397
70200000 | 9.4374 | -73.1597 | 70200000 | 13.824 | -132.6718
70340000 | 9.2879 | -72.943 | 70340000 | 13.389 | -132.4289
70480000 | 9.1116 | -72.7562 | 70480000 | 13.0363 | -132.1917
70620000 | 8.9577 | -72.5596 | 70620000 | 12.6551 | -131.9049
70760000 | 8.7827 | -72.3401 | 70760000 | 12.2964 | -131.5984
70900000 | 8.6123 | -72.1246 | 70900000 | 11.8768 | -131.3032
71040000 | 8.4607 | -71.9141 | 71040000 | 11.5082 | -131.037
71180000 | 8.3018 | -71.6977 | 71180000 | 11.1423 | -130.7127
71320000 | 8.1669 | -71.4689 | 71320000 | 10.8652 | -130.3891
71460000 | 8.0322 | -71.2288 | 71460000 | 10.451 | -130.0432
71600000 | 7.8644 | -71.006 | 71600000 | 10.0648 | -129.7317
71740000 | 7.7262 | -70.7608 | 71740000 | 9.7799 | -129.3299
71880000 | 7.5656 | -70.5288 | 71880000 | 9.4815 | -128.9752
72020000 | 7.4292 | -70.2801 | 72020000 | 9.1047 | -128.6163
72160000 | 7.3063 | -70.0454 | 72160000 | 8.7802 | -128.2623
72300000 | 7.1862 | -69.7878 | 72300000 | 8.4859 | -127.9001
72440000 | 7.035 | -69.5395 | 72440000 | 8.1445 | -127.5245
72580000 | 6.9213 | -69.3037 | 72580000 | 7.9275 | -127.121
72720000 | 6.81 | -69.0461 | 72720000 | 7.6273 | -126.7413
72860000 | 6.7004 | -68.7902 | 72860000 | 7.3105 | -126.334
73000000 | 6.5909 | -68.5497 | 73000000 | 7.0895 | -125.9262
73140000 | 6.4755 | -68.305 | 73140000 | 6.8403 | -125.5303
73280000 | 6.379 | -68.0417 | 73280000 | 6.5535 | -125.1286
73420000 | 6.2916 | -67.7892 | 73420000 | 6.329 | -124.738
73560000 | 6.176 | -67.5357 | 73560000 | 6.0688 | -124.3578
73700000 | 6.0759 | -67.3056 | 73700000 | 5.8579 | -123.9613
73840000 | 5.9815 | -67.0426 | 73840000 | 5.6576 | -123.5785
73980000 | 5.8886 | -66.7943 | 73980000 | 5.4451 | -123.1487
74120000 | 5.8085 | -66.5426 | 74120000 | 5.2451 | -122.7301
74260000 | 5.7267 | -66.2892 | 74260000 | 5.05 | -122.3982
74400000 | 5.6561 | -66.0375 | 74400000 | 4.8279 | -121.9772
74540000 | 5.564 | -65.7968 | 74540000 | 4.613 | -121.5766
74680000 | 5.5108 | -65.5566 | 74680000 | 4.478 | -121.1725
74820000 | 5.4264 | -65.3179 | 74820000 | 4.2983 | -120.8037
74960000 | 5.3498 | -65.061 | 74960000 | 4.1434 | -120.3732
75100000 | 5.2593 | -64.8122 | 75100000 | 3.9533 | -120.0219
75240000 | 5.2026 | -64.5661 | 75240000 | 3.795 | -119.6371
75380000 | 5.1465 | -64.3332 | 75380000 | 3.6396 | -119.2199
75520000 | 5.0991 | -64.0841 | 75520000 | 3.4919 | -118.8589
75660000 | 5.0239 | -63.8498 | 75660000 | 3.3262 | -118.4759
75800000 | 4.9817 | -63.6044 | 75800000 | 3.2373 | -118.0954
75940000 | 4.9164 | -63.3835 | 75940000 | 3.1225 | -117.7079
76080000 | 4.8725 | -63.1319 | 76080000 | 2.9619 | -117.3464
76220000 | 4.818 | -62.8879 | 76220000 | 2.8895 | -116.9401
76360000 | 4.7582 | -62.6687 | 76360000 | 2.7482 | -116.5611
76500000 | 4.7083 | -62.4156 | 76500000 | 2.645 | -116.2063
76640000 | 4.6598 | -62.1949 | 76640000 | 2.529 | -115.8277
76780000 | 4.6309 | -61.9582 | 76780000 | 2.39 | -115.4677
76920000 | 4.5761 | -61.7338 | 76920000 | 2.3535 | -115.1035
77060000 | 4.5366 | -61.4973 | 77060000 | 2.1812 | -114.7447
77200000 | 4.5005 | -61.2677 | 77200000 | 2.1136 | -114.3818
77340000 | 4.4634 | -61.0454 | 77340000 | 2.0288 | -114.0197
77480000 | 4.4142 | -60.8176 | 77480000 | 1.9412 | -113.6772
77620000 | 4.3781 | -60.5968 | 77620000 | 1.8634 | -113.3099
77760000 | 4.3542 | -60.3723 | 77760000 | 1.795 | -112.9601
77900000 | 4.3142 | -60.1582 | 77900000 | 1.7049 | -112.6355
78040000 | 4.2725 | -59.9388 | 78040000 | 1.6847 | -112.2714
78180000 | 4.2607 | -59.7079 | 78180000 | 1.5723 | -111.948
78320000 | 4.2358 | -59.4852 | 78320000 | 1.5327 | -111.5985
78460000 | 4.1883 | -59.2689 | 78460000 | 1.4065 | -111.2761
78600000 | 4.1706 | -59.0749 | 78600000 | 1.3756 | -110.9643
78740000 | 4.1473 | -58.847 | 78740000 | 1.3163 | -110.623
78880000 | 4.1273 | -58.6285 | 78880000 | 1.2901 | -110.3212
79020000 | 4.1055 | -58.4253 | 79020000 | 1.2469 | -109.9732
79160000 | 4.0931 | -58.2193 | 79160000 | 1.178 | -109.6829
79300000 | 4.0599 | -58.0159 | 79300000 | 1.1267 | -109.3467
79440000 | 4.0389 | -57.8214 | 79440000 | 1.0723 | -109.029
79580000 | 4.017 | -57.6063 | 79580000 | 0.99706 | -108.7164
79720000 | 3.9882 | -57.4085 | 79720000 | 0.9714 | -108.4215
79860000 | 3.9715 | -57.1977 | 79860000 | 0.94063 | -108.1275
80000000 | 3.9552 | -56.993 | 80000000 | 0.86389 | -107.8194

## Pluse Echo
20 dB attenuator in line with probe
Pulse settings
45 MHz
12 dB gain
+- 5 V Pulse
System gain offset by attenuator (~ 0dB system)
Insertion loss: |  | -40 | dB
Bandwidth |  | 41.11%
Scope Trace
Time | Signal
9.224e-06 | 0
9.2245e-06 | 0.0015625
9.225e-06 | 0.0015625
9.2255e-06 | 0.0015625
9.226e-06 | 0.003125
9.2265e-06 | 0.0015625
9.227e-06 | 0.0015625
9.2275e-06 | 0.0015625
9.228e-06 | 0.0015625
9.2285e-06 | 0.0015625
9.229e-06 | 0.0015625
9.2295e-06 | 0.0015625
9.23e-06 | 0.0015625
9.2305e-06 | -0.0015625
9.231e-06 | 0
9.2315e-06 | -0.0015625
9.232e-06 | -0.003125
9.2325e-06 | -0.0015625
9.233e-06 | -0.0015625
9.2335e-06 | -0.0015625
9.234e-06 | 0
9.2345e-06 | -0.0015625
9.235e-06 | 0
9.2355e-06 | -0.0015625
9.236e-06 | 0
9.2365e-06 | 0.0015625
9.237e-06 | 0.0015625
9.2375e-06 | 0.0015625
9.238e-06 | 0.0015625
9.2385e-06 | 0
9.239e-06 | -0.0015625
9.2395e-06 | -0.0015625
9.24e-06 | -0.0015625
9.2405e-06 | -0.003125
9.241e-06 | -0.003125
9.2415e-06 | -0.0015625
9.242e-06 | -0.003125
9.2425e-06 | -0.0015625
9.243e-06 | -0.0015625
9.2435e-06 | 0
9.244e-06 | 0
9.2445e-06 | -0.0015625
9.245e-06 | 0
9.2455e-06 | 0.0015625
9.246e-06 | 0.0015625
9.2465e-06 | -0.0015625
9.247e-06 | 0
9.2475e-06 | 0
9.248e-06 | -0.0015625
9.2485e-06 | -0.0015625
9.249e-06 | 0
9.2495e-06 | 0
9.25e-06 | 0
9.2505e-06 | -0.0015625
9.251e-06 | -0.0015625
9.2515e-06 | -0.0015625
9.252e-06 | 0
9.2525e-06 | 0.0015625
9.253e-06 | 0
9.2535e-06 | 0.0015625
9.254e-06 | 0.0015625
9.2545e-06 | 0
9.255e-06 | 0
9.2555e-06 | -0.003125
9.256e-06 | -0.0015625
9.2565e-06 | -0.0015625
9.257e-06 | 0
9.2575e-06 | 0
9.258e-06 | 0
9.2585e-06 | 0.0015625
9.259e-06 | -0.0015625
9.2595e-06 | -0.0015625
9.26e-06 | 0.0015625
9.2605e-06 | 0.0015625
9.261e-06 | 0
9.2615e-06 | 0
9.262e-06 | 0
9.2625e-06 | 0
9.263e-06 | -0.0015625
9.2635e-06 | -0.0015625
9.264e-06 | -0.0015625
9.2645e-06 | -0.003125
9.265e-06 | -0.0015625
9.2655e-06 | -0.0015625
9.266e-06 | -0.0015625
9.2665e-06 | -0.0015625
9.267e-06 | -0.0015625
9.2675e-06 | -0.0015625
9.268e-06 | -0.0015625
9.2685e-06 | 0
9.269e-06 | 0
9.2695e-06 | 0
9.27e-06 | 0.0015625
9.2705e-06 | -0.0015625
9.271e-06 | -0.0015625
9.2715e-06 | -0.0015625
9.272e-06 | -0.0015625
9.2725e-06 | -0.003125
9.273e-06 | -0.0015625
9.2735e-06 | -0.0015625
9.274e-06 | 0
9.2745e-06 | -0.0015625
9.275e-06 | -0.003125
9.2755e-06 | -0.0015625
9.276e-06 | 0
9.2765e-06 | -0.0015625
9.277e-06 | 0
9.2775e-06 | 0
9.278e-06 | -0.0015625
9.2785e-06 | 0
9.279e-06 | 0
9.2795e-06 | 0
9.28e-06 | 0
9.2805e-06 | 0
9.281e-06 | 0
9.2815e-06 | 0
9.282e-06 | 0.0015625
9.2825e-06 | 0.0015625
9.283e-06 | 0
9.2835e-06 | 0
9.284e-06 | 0.0015625
9.2845e-06 | 0
9.285e-06 | 0
9.2855e-06 | 0
9.286e-06 | 0
9.2865e-06 | -0.0015625
9.287e-06 | -0.003125
9.2875e-06 | -0.003125
9.288e-06 | -0.003125
9.2885e-06 | -0.003125
9.289e-06 | -0.003125
9.2895e-06 | -0.0015625
9.29e-06 | -0.0015625
9.2905e-06 | -0.0015625
9.291e-06 | -0.0015625
9.2915e-06 | 0
9.292e-06 | 0.0015625
9.2925e-06 | 0
9.293e-06 | 0.0015625
9.2935e-06 | 0.0015625
9.294e-06 | 0.0015625
9.2945e-06 | 0.0015625
9.295e-06 | 0
9.2955e-06 | 0
9.296e-06 | -0.0015625
9.2965e-06 | -0.0015625
9.297e-06 | -0.003125
9.2975e-06 | -0.003125
9.298e-06 | -0.003125
9.2985e-06 | -0.0015625
9.299e-06 | -0.003125
9.2995e-06 | -0.0015625
9.3e-06 | -0.0015625
9.3005e-06 | -0.0015625
9.301e-06 | -0.0015625
9.3015e-06 | 0
9.302e-06 | 0
9.3025e-06 | 0
9.303e-06 | -0.0015625
9.3035e-06 | -0.0015625
9.304e-06 | -0.0015625
9.3045e-06 | -0.003125
9.305e-06 | -0.0015625
9.3055e-06 | -0.0015625
9.306e-06 | 0
9.3065e-06 | -0.003125
9.307e-06 | -0.0015625
9.3075e-06 | -0.003125
9.308e-06 | -0.0015625
9.3085e-06 | -0.0015625
9.309e-06 | 0
9.3095e-06 | 0
9.31e-06 | 0
9.3105e-06 | -0.0015625
9.311e-06 | -0.0015625
9.3115e-06 | -0.0015625
9.312e-06 | -0.0015625
9.3125e-06 | -0.0015625
9.313e-06 | -0.003125
9.3135e-06 | -0.0015625
9.314e-06 | 0
9.3145e-06 | -0.0015625
9.315e-06 | 0
9.3155e-06 | 0
9.316e-06 | 0.0015625
9.3165e-06 | 0.0015625
9.317e-06 | 0
9.3175e-06 | 0
9.318e-06 | 0.0015625
9.3185e-06 | -0.0015625
9.319e-06 | -0.0015625
9.3195e-06 | -0.0015625
9.32e-06 | -0.0015625
9.3205e-06 | -0.0015625
9.321e-06 | -0.0015625
9.3215e-06 | -0.0015625
9.322e-06 | -0.0015625
9.3225e-06 | -0.003125
9.323e-06 | -0.003125
9.3235e-06 | -0.0015625
9.324e-06 | -0.0015625
9.3245e-06 | -0.0015625
9.325e-06 | -0.0015625
9.3255e-06 | 0
9.326e-06 | 0
9.3265e-06 | 0
9.327e-06 | -0.0015625
9.3275e-06 | 0
9.328e-06 | 0
9.3285e-06 | 0.0015625
9.329e-06 | 0
9.3295e-06 | 0.003125
9.33e-06 | 0.0015625
9.3305e-06 | 0.0015625
9.331e-06 | -0.0015625
9.3315e-06 | 0
9.332e-06 | 0
9.3325e-06 | -0.0015625
9.333e-06 | -0.0015625
9.3335e-06 | -0.0015625
9.334e-06 | -0.003125
9.3345e-06 | -0.0015625
9.335e-06 | -0.0015625
9.3355e-06 | -0.0015625
9.336e-06 | -0.0015625
9.3365e-06 | 0
9.337e-06 | 0
9.3375e-06 | 0.0015625
9.338e-06 | 0.0015625
9.3385e-06 | 0.0015625
9.339e-06 | 0.0015625
9.3395e-06 | 0.0015625
9.34e-06 | 0.0046875
9.3405e-06 | 0.003125
9.341e-06 | 0.0015625
9.3415e-06 | 0.003125
9.342e-06 | 0.0015625
9.3425e-06 | -0.0015625
9.343e-06 | -0.003125
9.3435e-06 | -0.0046875
9.344e-06 | -0.003125
9.3445e-06 | -0.00625
9.345e-06 | -0.003125
9.3455e-06 | -0.0046875
9.346e-06 | -0.003125
9.3465e-06 | 0
9.347e-06 | 0
9.3475e-06 | 0
9.348e-06 | 0.0015625
9.3485e-06 | 0
9.349e-06 | 0.003125
9.3495e-06 | 0.0046875
9.35e-06 | 0.0015625
9.3505e-06 | 0
9.351e-06 | 0.0015625
9.3515e-06 | 0.0015625
9.352e-06 | 0
9.3525e-06 | 0
9.353e-06 | 0
9.3535e-06 | -0.0015625
9.354e-06 | 0
9.3545e-06 | -0.0046875
9.355e-06 | -0.003125
9.3555e-06 | -0.0015625
9.356e-06 | -0.003125
9.3565e-06 | -0.0015625
9.357e-06 | -0.0015625
9.3575e-06 | 0
9.358e-06 | 0.0015625
9.3585e-06 | 0
9.359e-06 | 0.0015625
9.3595e-06 | 0.0015625
9.36e-06 | 0.003125
9.3605e-06 | 0.003125
9.361e-06 | 0.003125
9.3615e-06 | 0.003125
9.362e-06 | 0.003125
9.3625e-06 | 0.0015625
9.363e-06 | -0.0015625
9.3635e-06 | 0
9.364e-06 | -0.003125
9.3645e-06 | -0.003125
9.365e-06 | -0.0015625
9.3655e-06 | -0.003125
9.366e-06 | -0.003125
9.3665e-06 | -0.0015625
9.367e-06 | -0.0015625
9.3675e-06 | -0.0015625
9.368e-06 | -0.0015625
9.3685e-06 | 0
9.369e-06 | 0
9.3695e-06 | 0.003125
9.37e-06 | 0.003125
9.3705e-06 | 0.003125
9.371e-06 | 0.003125
9.3715e-06 | 0.003125
9.372e-06 | 0.003125
9.3725e-06 | 0
9.373e-06 | 0
9.3735e-06 | 0
9.374e-06 | -0.003125
9.3745e-06 | -0.003125
9.375e-06 | -0.003125
9.3755e-06 | -0.0046875
9.376e-06 | -0.003125
9.3765e-06 | -0.0046875
9.377e-06 | -0.0046875
9.3775e-06 | -0.003125
9.378e-06 | -0.003125
9.3785e-06 | 0
9.379e-06 | -0.0015625
9.3795e-06 | 0
9.38e-06 | 0.003125
9.3805e-06 | 0.0015625
9.381e-06 | 0.003125
9.3815e-06 | 0.003125
9.382e-06 | 0.003125
9.3825e-06 | 0
9.383e-06 | 0.0015625
9.3835e-06 | 0
9.384e-06 | -0.0015625
9.3845e-06 | -0.0015625
9.385e-06 | -0.0015625
9.3855e-06 | -0.003125
9.386e-06 | -0.003125
9.3865e-06 | -0.003125
9.387e-06 | -0.0046875
9.3875e-06 | -0.0015625
9.388e-06 | -0.003125
9.3885e-06 | -0.0015625
9.389e-06 | 0
9.3895e-06 | 0.0015625
9.39e-06 | 0.003125
9.3905e-06 | 0.0015625
9.391e-06 | 0.0015625
9.3915e-06 | 0.003125
9.392e-06 | 0.0046875
9.3925e-06 | 0.003125
9.393e-06 | 0.0015625
9.3935e-06 | 0.003125
9.394e-06 | 0.0015625
9.3945e-06 | -0.0015625
9.395e-06 | 0
9.3955e-06 | 0.0015625
9.396e-06 | 0
9.3965e-06 | -0.0015625
9.397e-06 | -0.0015625
9.3975e-06 | -0.0015625
9.398e-06 | -0.0015625
9.3985e-06 | -0.0015625
9.399e-06 | -0.0015625
9.3995e-06 | -0.0015625
9.4e-06 | 0
9.4005e-06 | 0
9.401e-06 | 0
9.4015e-06 | 0
9.402e-06 | 0
9.4025e-06 | 0.0015625
9.403e-06 | 0.0015625
9.4035e-06 | 0.0015625
9.404e-06 | 0.0015625
9.4045e-06 | 0
9.405e-06 | 0.0015625
9.4055e-06 | 0.0015625
9.406e-06 | 0
9.4065e-06 | 0
9.407e-06 | -0.003125
9.4075e-06 | -0.003125
9.408e-06 | -0.0015625
9.4085e-06 | -0.0046875
9.409e-06 | -0.0046875
9.4095e-06 | -0.003125
9.41e-06 | -0.0015625
9.4105e-06 | -0.0015625
9.411e-06 | 0
9.4115e-06 | 0
9.412e-06 | 0.0015625
9.4125e-06 | 0.0015625
9.413e-06 | 0.0046875
9.4135e-06 | 0.003125
9.414e-06 | 0.003125
9.4145e-06 | 0.003125
9.415e-06 | 0.003125
9.4155e-06 | 0.0015625
9.416e-06 | 0
9.4165e-06 | 0.0015625
9.417e-06 | 0
9.4175e-06 | 0.0015625
9.418e-06 | 0
9.4185e-06 | -0.0015625
9.419e-06 | -0.0015625
9.4195e-06 | -0.0015625
9.42e-06 | -0.0015625
9.4205e-06 | -0.0015625
9.421e-06 | 0
9.4215e-06 | 0
9.422e-06 | 0
9.4225e-06 | 0
9.423e-06 | 0
9.4235e-06 | 0.0015625
9.424e-06 | 0
9.4245e-06 | 0.003125
9.425e-06 | 0.0015625
9.4255e-06 | 0.003125
9.426e-06 | 0.003125
9.4265e-06 | 0.0015625
9.427e-06 | 0
9.4275e-06 | 0
9.428e-06 | 0.0015625
9.4285e-06 | 0
9.429e-06 | 0.0015625
9.4295e-06 | 0.0015625
9.43e-06 | 0.003125
9.4305e-06 | 0
9.431e-06 | 0
9.4315e-06 | 0
9.432e-06 | 0
9.4325e-06 | 0.0015625
9.433e-06 | 0.0015625
9.4335e-06 | 0.003125
9.434e-06 | 0.0046875
9.4345e-06 | 0.0046875
9.435e-06 | 0.0046875
9.4355e-06 | 0.00625
9.436e-06 | 0.00625
9.4365e-06 | 0.00625
9.437e-06 | 0.0046875
9.4375e-06 | 0.003125
9.438e-06 | 0.0046875
9.4385e-06 | 0
9.439e-06 | -0.0015625
9.4395e-06 | -0.003125
9.44e-06 | -0.0046875
9.4405e-06 | -0.0078125
9.441e-06 | -0.009375
9.4415e-06 | -0.0125
9.442e-06 | -0.0125
9.4425e-06 | -0.015625
9.443e-06 | -0.017188
9.4435e-06 | -0.01875
9.444e-06 | -0.01875
9.4445e-06 | -0.01875
9.445e-06 | -0.01875
9.4455e-06 | -0.020313
9.446e-06 | -0.01875
9.4465e-06 | -0.01875
9.447e-06 | -0.020313
9.4475e-06 | -0.020313
9.448e-06 | -0.01875
9.4485e-06 | -0.01875
9.449e-06 | -0.017188
9.4495e-06 | -0.017188
9.45e-06 | -0.015625
9.4505e-06 | -0.014063
9.451e-06 | -0.010937
9.4515e-06 | -0.0078125
9.452e-06 | -0.0046875
9.4525e-06 | -0.0015625
9.453e-06 | 0.003125
9.4535e-06 | 0.0078125
9.454e-06 | 0.0125
9.4545e-06 | 0.017188
9.455e-06 | 0.021875
9.4555e-06 | 0.029687
9.456e-06 | 0.032813
9.4565e-06 | 0.0375
9.457e-06 | 0.04375
9.4575e-06 | 0.045312
9.458e-06 | 0.048438
9.4585e-06 | 0.05
9.459e-06 | 0.048438
9.4595e-06 | 0.046875
9.46e-06 | 0.04375
9.4605e-06 | 0.042188
9.461e-06 | 0.035937
9.4615e-06 | 0.032813
9.462e-06 | 0.025
9.4625e-06 | 0.01875
9.463e-06 | 0.014063
9.4635e-06 | 0.0078125
9.464e-06 | -0.0015625
9.4645e-06 | -0.00625
9.465e-06 | -0.0125
9.4655e-06 | -0.020313
9.466e-06 | -0.023438
9.4665e-06 | -0.029687
9.467e-06 | -0.034375
9.4675e-06 | -0.039062
9.468e-06 | -0.042188
9.4685e-06 | -0.045312
9.469e-06 | -0.048438
9.4695e-06 | -0.05
9.47e-06 | -0.051562
9.4705e-06 | -0.051562
9.471e-06 | -0.05
9.4715e-06 | -0.048438
9.472e-06 | -0.046875
9.4725e-06 | -0.040625
9.473e-06 | -0.035937
9.4735e-06 | -0.029687
9.474e-06 | -0.021875
9.4745e-06 | -0.014063
9.475e-06 | -0.00625
9.4755e-06 | 0
9.476e-06 | 0.009375
9.4765e-06 | 0.017188
9.477e-06 | 0.023438
9.4775e-06 | 0.03125
9.478e-06 | 0.0375
9.4785e-06 | 0.040625
9.479e-06 | 0.045312
9.4795e-06 | 0.046875
9.48e-06 | 0.048438
9.4805e-06 | 0.046875
9.481e-06 | 0.045312
9.4815e-06 | 0.042188
9.482e-06 | 0.039062
9.4825e-06 | 0.034375
9.483e-06 | 0.029687
9.4835e-06 | 0.023438
9.484e-06 | 0.017188
9.4845e-06 | 0.010937
9.485e-06 | 0.00625
9.4855e-06 | 0.0015625
9.486e-06 | -0.0046875
9.4865e-06 | -0.010937
9.487e-06 | -0.015625
9.4875e-06 | -0.020313
9.488e-06 | -0.023438
9.4885e-06 | -0.026562
9.489e-06 | -0.028125
9.4895e-06 | -0.03125
9.49e-06 | -0.03125
9.4905e-06 | -0.032813
9.491e-06 | -0.03125
9.4915e-06 | -0.029687
9.492e-06 | -0.028125
9.4925e-06 | -0.025
9.493e-06 | -0.020313
9.4935e-06 | -0.015625
9.494e-06 | -0.0125
9.4945e-06 | -0.0078125
9.495e-06 | -0.0046875
9.4955e-06 | 0
9.496e-06 | 0.0046875
9.4965e-06 | 0.0078125
9.497e-06 | 0.0125
9.4975e-06 | 0.015625
9.498e-06 | 0.01875
9.4985e-06 | 0.020313
9.499e-06 | 0.021875
9.4995e-06 | 0.023438
9.5e-06 | 0.021875
9.5005e-06 | 0.021875
9.501e-06 | 0.021875
9.5015e-06 | 0.020313
9.502e-06 | 0.017188
9.5025e-06 | 0.014063
9.503e-06 | 0.009375
9.5035e-06 | 0.0046875
9.504e-06 | 0.0015625
9.5045e-06 | -0.003125
9.505e-06 | -0.0046875
9.5055e-06 | -0.0078125
9.506e-06 | -0.010937
9.5065e-06 | -0.0125
9.507e-06 | -0.017188
9.5075e-06 | -0.01875
9.508e-06 | -0.01875
9.5085e-06 | -0.020313
9.509e-06 | -0.021875
9.5095e-06 | -0.020313
9.51e-06 | -0.020313
9.5105e-06 | -0.01875
9.511e-06 | -0.015625
9.5115e-06 | -0.0125
9.512e-06 | -0.010937
9.5125e-06 | -0.0078125
9.513e-06 | -0.0046875
9.5135e-06 | -0.0015625
9.514e-06 | 0.0015625
9.5145e-06 | 0.003125
9.515e-06 | 0.0078125
9.5155e-06 | 0.009375
9.516e-06 | 0.0125
9.5165e-06 | 0.014063
9.517e-06 | 0.015625
9.5175e-06 | 0.017188
9.518e-06 | 0.017188
9.5185e-06 | 0.017188
9.519e-06 | 0.017188
9.5195e-06 | 0.015625
9.52e-06 | 0.014063
9.5205e-06 | 0.0125
9.521e-06 | 0.0125
9.5215e-06 | 0.009375
9.522e-06 | 0.0078125
9.5225e-06 | 0.00625
9.523e-06 | 0.003125
9.5235e-06 | 0
9.524e-06 | 0
9.5245e-06 | -0.0015625
9.525e-06 | -0.003125
9.5255e-06 | -0.003125
9.526e-06 | -0.0046875
9.5265e-06 | -0.0078125
9.527e-06 | -0.009375
9.5275e-06 | -0.010937
9.528e-06 | -0.0125
9.5285e-06 | -0.0125
9.529e-06 | -0.014063
9.5295e-06 | -0.015625
9.53e-06 | -0.0125
9.5305e-06 | -0.0125
9.531e-06 | -0.010937
9.5315e-06 | -0.009375
9.532e-06 | -0.0078125
9.5325e-06 | -0.00625
9.533e-06 | -0.003125
9.5335e-06 | 0
9.534e-06 | 0.0015625
9.5345e-06 | 0
9.535e-06 | 0.0015625
9.5355e-06 | 0.003125
9.536e-06 | 0.0046875
9.5365e-06 | 0.00625
9.537e-06 | 0.00625
9.5375e-06 | 0.0046875
9.538e-06 | 0.00625
9.5385e-06 | 0.00625
9.539e-06 | 0.0046875
9.5395e-06 | 0.0046875
9.54e-06 | 0.003125
9.5405e-06 | 0.003125
9.541e-06 | 0.003125
9.5415e-06 | 0.003125
9.542e-06 | 0.003125
9.5425e-06 | 0
9.543e-06 | 0.0015625
9.5435e-06 | 0
9.544e-06 | 0
9.5445e-06 | 0.0015625
9.545e-06 | 0
9.5455e-06 | 0
9.546e-06 | 0.0015625
9.5465e-06 | -0.0015625
9.547e-06 | 0
9.5475e-06 | -0.0015625
9.548e-06 | 0
9.5485e-06 | 0
9.549e-06 | -0.0015625
9.5495e-06 | -0.0015625
9.55e-06 | -0.0015625
9.5505e-06 | -0.0015625
9.551e-06 | -0.0046875
9.5515e-06 | -0.003125
9.552e-06 | -0.0015625
9.5525e-06 | -0.0015625
9.553e-06 | -0.0015625
9.5535e-06 | 0
9.554e-06 | 0.0015625
9.5545e-06 | 0.0015625
9.555e-06 | 0.0015625
9.5555e-06 | 0.003125
9.556e-06 | 0.0046875
9.5565e-06 | 0.003125
9.557e-06 | 0.00625
9.5575e-06 | 0.003125
9.558e-06 | 0.003125
9.5585e-06 | 0.0015625
9.559e-06 | 0.0015625
9.5595e-06 | 0
9.56e-06 | -0.0015625
9.5605e-06 | -0.003125
9.561e-06 | -0.003125
9.5615e-06 | -0.0046875
9.562e-06 | -0.00625
9.5625e-06 | -0.0046875
9.563e-06 | -0.00625
9.5635e-06 | -0.0046875
9.564e-06 | -0.003125
9.5645e-06 | -0.003125
9.565e-06 | -0.0015625
9.5655e-06 | 0
9.566e-06 | 0.0015625
9.5665e-06 | 0
9.567e-06 | 0.0015625
9.5675e-06 | 0.0015625
9.568e-06 | 0
9.5685e-06 | 0.0015625
9.569e-06 | 0.0015625
9.5695e-06 | 0.003125
9.57e-06 | 0.0015625
9.5705e-06 | 0.0015625
9.571e-06 | 0
9.5715e-06 | 0.0015625
9.572e-06 | 0.003125
9.5725e-06 | 0.003125
9.573e-06 | 0.003125
9.5735e-06 | 0.003125
9.574e-06 | 0.003125
9.5745e-06 | 0.003125
9.575e-06 | 0.0015625
9.5755e-06 | 0.0015625
9.576e-06 | 0.003125
9.5765e-06 | 0.003125
9.577e-06 | 0.003125
9.5775e-06 | 0.0015625
9.578e-06 | 0.003125
9.5785e-06 | 0.0015625
9.579e-06 | 0
9.5795e-06 | 0.0015625
9.58e-06 | -0.0015625
9.5805e-06 | 0
9.581e-06 | -0.0015625
9.5815e-06 | -0.003125
9.582e-06 | -0.003125
9.5825e-06 | -0.003125
9.583e-06 | -0.003125
9.5835e-06 | -0.00625
9.584e-06 | -0.003125
9.5845e-06 | -0.00625
9.585e-06 | -0.0046875
9.5855e-06 | -0.003125
9.586e-06 | -0.003125
9.5865e-06 | -0.003125
9.587e-06 | -0.0015625
9.5875e-06 | 0
9.588e-06 | 0
9.5885e-06 | 0
9.589e-06 | 0.0015625
9.5895e-06 | 0.0015625
9.59e-06 | 0.003125
9.5905e-06 | 0.003125
9.591e-06 | 0.0015625
9.5915e-06 | 0.0015625
9.592e-06 | 0.003125
9.5925e-06 | 0.0015625
9.593e-06 | 0.0015625
9.5935e-06 | 0.0015625
9.594e-06 | 0.0015625
9.5945e-06 | 0.0015625
9.595e-06 | 0
9.5955e-06 | 0
9.596e-06 | 0.0015625
9.5965e-06 | 0.0015625
9.597e-06 | 0.003125
9.5975e-06 | 0.0015625
9.598e-06 | 0.003125
9.5985e-06 | 0.0015625
9.599e-06 | 0.0015625
9.5995e-06 | 0.0015625
9.6e-06 | 0
9.6005e-06 | 0
9.601e-06 | 0.0015625
9.6015e-06 | 0.0015625
9.602e-06 | 0.0015625
9.6025e-06 | -0.0015625
9.603e-06 | -0.0015625
9.6035e-06 | 0
9.604e-06 | -0.0015625
9.6045e-06 | -0.0015625
9.605e-06 | -0.003125
9.6055e-06 | -0.0015625
9.606e-06 | -0.0015625
9.6065e-06 | -0.0015625
9.607e-06 | -0.003125
9.6075e-06 | -0.003125
9.608e-06 | -0.003125
9.6085e-06 | -0.0015625
9.609e-06 | -0.0015625
9.6095e-06 | 0
9.61e-06 | 0.0015625
9.6105e-06 | 0.0015625
9.611e-06 | 0
9.6115e-06 | 0
9.612e-06 | 0.0015625
9.6125e-06 | 0
9.613e-06 | 0.0015625
9.6135e-06 | 0.0015625
9.614e-06 | 0.003125
9.6145e-06 | 0
9.615e-06 | 0
9.6155e-06 | -0.0015625
9.616e-06 | 0
9.6165e-06 | 0
9.617e-06 | 0
9.6175e-06 | 0
9.618e-06 | 0
9.6185e-06 | 0
9.619e-06 | 0
9.6195e-06 | 0
9.62e-06 | 0
9.6205e-06 | 0.0015625
9.621e-06 | 0
9.6215e-06 | 0
9.622e-06 | 0.0015625
9.6225e-06 | 0
9.623e-06 | 0
9.6235e-06 | 0
9.624e-06 | 0
9.6245e-06 | -0.0015625
9.625e-06 | -0.0015625
9.6255e-06 | -0.0015625
9.626e-06 | -0.0015625
9.6265e-06 | -0.0015625
9.627e-06 | -0.0015625
9.6275e-06 | 0
9.628e-06 | 0
9.6285e-06 | -0.0015625
9.629e-06 | 0
9.6295e-06 | 0
9.63e-06 | 0
9.6305e-06 | -0.0015625
9.631e-06 | -0.0015625
9.6315e-06 | 0
9.632e-06 | 0
9.6325e-06 | 0
9.633e-06 | 0
9.6335e-06 | 0.0015625
9.634e-06 | 0.0015625
9.6345e-06 | 0
9.635e-06 | 0
9.6355e-06 | 0
9.636e-06 | 0.0015625
9.6365e-06 | 0
9.637e-06 | 0
9.6375e-06 | -0.0015625
9.638e-06 | -0.0015625
9.6385e-06 | 0
9.639e-06 | -0.0015625
9.6395e-06 | -0.0046875
9.64e-06 | -0.003125
9.6405e-06 | -0.0015625
9.641e-06 | -0.003125
9.6415e-06 | -0.0015625
9.642e-06 | 0
9.6425e-06 | -0.0015625
9.643e-06 | 0.0015625
9.6435e-06 | 0
9.644e-06 | 0.0015625
9.6445e-06 | 0.0015625
9.645e-06 | 0
9.6455e-06 | 0
9.646e-06 | -0.0015625
9.6465e-06 | -0.0015625
9.647e-06 | -0.003125
9.6475e-06 | -0.003125
9.648e-06 | -0.0046875
9.6485e-06 | -0.003125
9.649e-06 | -0.003125
9.6495e-06 | -0.003125
9.65e-06 | -0.0015625
9.6505e-06 | -0.0015625
9.651e-06 | 0
9.6515e-06 | 0
9.652e-06 | 0.0015625
9.6525e-06 | 0.0015625
9.653e-06 | 0.003125
9.6535e-06 | 0.003125
9.654e-06 | 0.003125
9.6545e-06 | 0.0015625
9.655e-06 | 0.0015625
9.6555e-06 | 0
9.656e-06 | 0
9.6565e-06 | 0
9.657e-06 | 0
9.6575e-06 | 0
9.658e-06 | 0
9.6585e-06 | -0.0015625
9.659e-06 | 0
9.6595e-06 | 0
9.66e-06 | 0
9.6605e-06 | 0.0015625
9.661e-06 | 0.0015625
9.6615e-06 | 0.003125
9.662e-06 | 0.0015625
9.6625e-06 | 0
9.663e-06 | 0
9.6635e-06 | 0
9.664e-06 | 0
9.6645e-06 | 0
9.665e-06 | 0
9.6655e-06 | -0.0015625
9.666e-06 | -0.0015625
9.6665e-06 | 0
9.667e-06 | -0.003125
9.6675e-06 | -0.0015625
9.668e-06 | -0.003125
9.6685e-06 | -0.0015625
9.669e-06 | 0
9.6695e-06 | 0
9.67e-06 | -0.0015625
9.6705e-06 | 0
9.671e-06 | -0.0015625
9.6715e-06 | -0.0015625
9.672e-06 | -0.0015625
9.6725e-06 | -0.0015625
9.673e-06 | 0
9.6735e-06 | -0.0015625
9.674e-06 | 0
9.6745e-06 | -0.0015625
9.675e-06 | -0.0015625
9.6755e-06 | 0
9.676e-06 | 0.0015625
9.6765e-06 | 0
9.677e-06 | 0.0015625
9.6775e-06 | 0
9.678e-06 | 0.0015625
9.6785e-06 | -0.0015625
9.679e-06 | -0.0015625
9.6795e-06 | -0.0015625
9.68e-06 | 0
9.6805e-06 | 0
9.681e-06 | 0
9.6815e-06 | 0
9.682e-06 | 0.0015625
9.6825e-06 | 0.0015625
9.683e-06 | 0
9.6835e-06 | 0.0015625
9.684e-06 | 0.003125
9.6845e-06 | 0.003125
9.685e-06 | 0.003125
9.6855e-06 | 0.003125
9.686e-06 | 0.0015625
9.6865e-06 | 0
9.687e-06 | 0
9.6875e-06 | 0
9.688e-06 | -0.0015625
9.6885e-06 | -0.0015625
9.689e-06 | -0.0015625
9.6895e-06 | -0.0015625
9.69e-06 | 0
9.6905e-06 | -0.0015625
9.691e-06 | 0
9.6915e-06 | -0.0015625
9.692e-06 | 0
9.6925e-06 | 0
9.693e-06 | 0.0015625
9.6935e-06 | 0.0015625
9.694e-06 | 0.003125
9.6945e-06 | 0
9.695e-06 | 0.0015625
9.6955e-06 | 0
9.696e-06 | -0.0015625
9.6965e-06 | 0
9.697e-06 | -0.0015625
9.6975e-06 | -0.0015625
9.698e-06 | -0.0015625
9.6985e-06 | -0.0046875
9.699e-06 | -0.0015625
9.6995e-06 | 0
9.7e-06 | 0
9.7005e-06 | 0
9.701e-06 | 0.0015625
9.7015e-06 | 0
9.702e-06 | 0.0015625
9.7025e-06 | 0
9.703e-06 | 0
9.7035e-06 | 0
9.704e-06 | 0
9.7045e-06 | 0
9.705e-06 | 0
9.7055e-06 | 0.0015625
9.706e-06 | 0
9.7065e-06 | -0.0015625
9.707e-06 | -0.0015625
9.7075e-06 | 0
9.708e-06 | 0
9.7085e-06 | -0.0015625
9.709e-06 | -0.0015625
9.7095e-06 | 0
9.71e-06 | 0
9.7105e-06 | 0
9.711e-06 | -0.0015625
9.7115e-06 | -0.0015625
9.712e-06 | -0.0015625
9.7125e-06 | -0.0015625
9.713e-06 | 0
9.7135e-06 | 0.003125
9.714e-06 | 0.0015625
9.7145e-06 | 0.0015625
9.715e-06 | 0.0015625
9.7155e-06 | 0.0015625
9.716e-06 | 0.003125
9.7165e-06 | 0.0015625
9.717e-06 | 0.0015625
9.7175e-06 | 0
9.718e-06 | 0.0015625
9.7185e-06 | -0.0015625
9.719e-06 | -0.0015625
9.7195e-06 | -0.0015625
9.72e-06 | -0.0015625
9.7205e-06 | -0.003125
9.721e-06 | -0.003125
9.7215e-06 | -0.003125
9.722e-06 | -0.0015625
9.7225e-06 | -0.0015625
9.723e-06 | -0.0015625
9.7235e-06 | -0.0015625